## Slide 1: Design Challenges for Secure Implantable Medical Devices
Wayne Burleson
Department of Electrical and Computer Engineering
Shane Clark, Ben Ransford, Kevin Fu, Department of Computer Science
Physical Layer Security
This material is based upon work supported by: the Armstrong Fund for Science; the National Science Foundation under Grants No. 831244, 0923313 and 0964641; Cooperative Agreement No. 90TR0003/01 from the Department of Health and Human Services; two NSF Graduate Research Fellowships; and a Sloan Research Fellowship. Its contents are solely the responsibility of the authors and do not necessarily represent the official views of DHHS or NSF.
University of Massachusetts Amherst
burleson@ecs.umass.edu

## Slide 2: Implantable and Wearable Medical Devices
Bio-Medical
  EEG Electroencephalography
  ECG Electrocardiogram
  EMG Electromyography (muscular)
  Blood pressure
  Blood SpO2
  Blood pH
  Glucose sensor
  Respiration
  Temperature
  Fall detection
  Ocular/cochlear prosthesis
  Digestive tract tracking
  Digestive tract imaging
Sports performance
  Distance
  Speed
  Posture (Body Position)
  Sports training aid
Cyber-human interfaces
Body Area
Network (BAN)
Images courtesy CSEM , Switzerland

## Slide 3: Security and Privacy in Implantable Medical Devices
IMD’s are an increasingly important technology
  Leveraging many recent technologies in Nano/Bio/Info
  Possible solutions to major societal problems
    Clinical
    Research
  Many types of IMDs (see taxonomy coming up)
Security and Privacy increasingly relevant in modern society
  Fundamental human rights
  Quality of life, Related to safety/health
  Acceptance of new technologies
Combining 1. and 2., IMD Security and Privacy involves:
    Protecting human life, health and well-being
    Protecting health information and record privacy
    Engineering Challenges!

## Slide 4: IMD Examples
Existing
  Glucose sensor and insulin pump
  Pacemaker/defibrillator
  Neurostimulator
  Cochlear implant
Emerging
  Ingestible “smart-pills”
  Drug delivery
  Sub-cutaneous biosensor
  Brain implant
  Deep cardiac implant
  Smart Orthodontia
  Glaucoma sensors and ocular implants
Futuristic
  Body 2.0 - Continuous Monitoring of the Human Body
  Bio-reactors
  Cyber-human Interfaces
concept illustration from yankodesign
Smart pill - Proteus biomedical
Pacemaker - Medtronic
Subcutaneous biosensor – EPFL-Nanotera
Neurostimulator
Cochlear implant

## Slide 5: Axes for a taxonomy of IMDs
Physical location/depth, procedure, lifetime,
Sensing/Actuating functions, (sense, deliver drugs or stimulus, grow tissue!)
Computational capabilities
Data storage
Communication: bandwidth, up-link, down-link, inter-device? Positioning system (IPS), distance to reader, noise
Energy requirements, (memory, communication, computation,) powering, harvesting, storage, (battery or capacitive)?
Vulnerabilities. Security functions (access control, authentication, encryption)
Reliability and Failure modes

## Slide 6: Security Goals for IMD Design
Incorporate security early.
Encrypt sensitive traffic.
Authenticate third-party devices.
Use well-studied cryptographic building blocks.
Do not rely on security through obscurity.
Use industry-standard source-code analysis.
Develop a realistic threat model.

## Slide 7: Threat model – Understand your adversary!
Motives:
  Identity Theft
  Violence
  Insurance fraud
  Counterfeit devices
Resources:
  Individual
  Organization
  Nation-state…
Attack vectors:
  Wireless interfaces (eavesdropping, jamming, man-in-middle)
  Data/control from unauthenticated sources
  Data retention in discarded devices

## Slide 8: Pacemakers, Defibrillators (UM Amherst, Harvard, Beth Israel)
Many medical devices rely on wireless connectivity for remote monitoring, remote therapies and software updates.
Recent research identified several attacks and defenses for implantable cardiac defibrillators
  Wireless communications were unencrypted and unauthenticated
  Leading to several lethal vulnerabilities
Extensions to numerous other emerging implantable devices
Pacemakers and Implantable Cardiac Defibrillators: Software Radio Attacks and Zero-Power Defenses. D. Halperin, T. Heydt-Benjamin, B. Ransford, S. Clark, B. Defend, W. Morgan, K. Fu, T. Kohno, and W. Maisel. In Proceedings of the 29th Annual IEEE Symposium on Security and Privacy, May 2008. Best Paper Award
TR 35

## Slide 9: Benefits of Wireless
Easier communication with implant
Remote monitoring
Notes: Having wireless in these devices has multiple benefits.
First, it makes it easier to communicate with the implant itself.
More importantly, it enables remote monitoring of the patient’s status.
The patient can be in her home; she would have an implant reader that periodically queries the implant for patient’s vital signals. The reader then forwards this information to the medical professionals over the internet.

## Slide 10: Benefits of Wireless
Easier communication with implant
Remote monitoring
Reduces hospital visits by 40% and cost per visit by $1800
[Journal of the American College of Cardiology, 2011]
What about security?
Notes: Studies show that such remote monitoring reduces the number of hospital visits by 40% and also reduces the cost of each clinical visit by about 1800 dollars.
So having wireless in medical implants has many benefits. But, it also brings along new security and privacy issues.

## Slide 11: Security Attacks
1) Passive attack: Eavesdrop on private data
Patient diagnosis, vital signs
2) Active attack: Send unauthorized commands
Turn off therapies,
[Halperin’08] demonstrated attacks using software radios
deliver electric shock
Notes: In particular there are two attacks that are important.
First a passive attack. In this case, the attacker listens on the medium. Whenever the implant transmits, the attacker snoops on these transmission to obtain the patient’s private data. This could be the patient’s diagnosis or vital signal like EKG. Such a snooping is a significant violation of the patient’s privacy.
Second an active attack. In this case the attacker sends unauthorized commands to the implant.
For example, the attacker here can send wireless commands which turn off therapies on the implant.
These passive and active attacks are not just fantasies. They have been demonstrated in practice using off-the-shelf software radios. For example, the researchers have showed that an attacker can make a cardiac defibrillators deliver an electric shock to the patient’s heart.

## Slide 12: Insulin Pump Systems
Patient-controlled open-loop systems used to monitor and stabilize glucose levels.
Several researchers have highlighted security and privacy risks in insulin pump systems.
    Wireless forgery of insulin readings
    Wireless administration and potentially fatal over-dosage.
C. Li, A. Raghunathan, and N. K. Jha. Hijacking an insulin
pump: Security attacks and defenses for a diabetes therapy
system. In Proceedings of the 13th IEEE International
Conference on e-Health Networking, Applications, and
Services, Healthcom ’11, June 2011.
N. Paul, T. Kohno, and D. C. Klonoff. A review of the
security of insulin pump infusion systems. Journal of
Diabetes Science and Technology, 5(6):1557–1562,
November 2011.

## Slide 13: Cross-cutting Concerns
When and how to apply encryption
  Authentication and Key management
  Lightweight ciphers (stream and block)
  Physical layer security
  Appropriate failure modes
Novel approaches to authentication
  Ultrasonic distance-bounding
  Auxiliary “helper” devices
  PUFs
Cyber-human systems
  Human on both ends of the system
    Controlling
    Sensing
  Humans in the loop

## Slide 14: Lightweight Cryptography for Bio-sensors
Hummingbird Stream Cipher
Glucose sensor
AES Block Cipher
Ocular implant
S. Guan, J. Gu, Z. Shen, J. Wang, Y. Huang, and A. Mason.
A wireless powered implantable bio-sensor tag system-on-chip for continuous glucose monitoring. BioCAS 2011.
C. Beck, D. Masny, W. Geiselmann, and G. Bretthauer.
Block cipher based security for severely resource-constrained implantable medical devices. International Symposium on Applied Sciences in Biomedical and Communication Technologies, ISABEL 2011.

## Slide 15: External “protector devices”
Sorber et al (Dartmouth), An Amulet for trustworthy wearable mHealth, HotMobile 2012

## Slide 16: Protecting existing IMDs
Gollakota et al (MIT, UMASS), They Can Hear Your Heartbeats: Non-Invasive Security for Implanted Medical Devices, SIGCOMM 2011 (Best Paper)

## Slide 17: Secure Platform for Bio-sensing (Umass, EPFL, Bochum)
Implanted Devices
Disposable Diagnostic
Applications
  Disposable Diagnostic
    Low-cost, infectious disease detection (malaria, HIV, dengue, cholera)
  Implantable Device
    Sub-cutaneous multi-function sensor (drugs, antibodies)
    Glucose/Lactate in Trauma victims
Security Technology
  NFC Cell Phone
  EPC Class 1, Gen 2 protocol
  PRESENT Block Cipher (Encryption, Signing, Authentication)
  PUF for low-cost ID and Challenge-Response
Images: Disposable Diagnostic: Gentag.com,
Sub-cutaneous Implant: LSI, EPFL, NanoTera
2-element biochip: CBBB, Clemson University

## Slide 18: Power/Energy Challenges
Remote powered systems (RFID) limited to 10’s of microwatts
Near field powering improves this to milliwatts
Current energy harvesting systems similarly limited…
Small batteries typically store several 1000 Joules.
Over several years of operation, this translates to 10’s of microwatts
Batteries are still large and heavy
Rechargable batteries dissipate
heat and have safety concerns
Non-rechargeable batteries
require surgery for replacement
Brain implants can not incur more than 1 degree Celsius temperature gradient without safety concerns

## Slide 19: Design Tension Challenges
Safety/Utility goals
Data access
Data accuracy
Device identification
Configurability
Updatable software
Multi-device coordination
Auditable
Resource efficient
Security/Privacy goals
Authorization (personal, role-based, IMD selection)
Availability
Device software and settings
Device-existence privacy
Device-type privacy
Specific-device ID privacy
Measurement and Log Privacy
Bearer privacy
Data integrity
From D. Halperin et al, “Security and Privacy for Implantable Medical Devices”, IEEE Pervasive Computing, 2008

## Slide 20: Design for Medical is different!
“Medical marches to a different cadence than most of the electronics industry. Design cycles can stretch from three to five years and cost $10-15 million, thanks to the lengthy regulatory process. The product lifecycles can also extend over a 20 year time span.”
Boston Scientific
  What is the role of FDA and other regulators?
  - FDA currently regulates safety, but not security

## Slide 21: Trends in VLSI Research
Design Challenges
  Area
  Performance
  Complexity
  Test/Yield
  Power
  Flexibility
  Reliability
    Process
    Voltage
    Temperature
  Security/Privacy
1970’s
1980’s
1990’s
2000’s
2010’s
Driving Applications
  Microprocessors
  DSP
  Video
  Wireless
  Hand-sets
  Smart Cards
  Sensor Networks
  RFID
  Medical Devices
  Smart Dust
  …

## Slide 22
(co-located with IEEE ISMICT in nearby Montreux, Switzerland, www.ismict2011.org)
Speakers:
K. Fu Umass Amherst, USA
S. Capkun, ETHZ, CH
S. Carrara, EPFL, CH
J. Huiskens, IMEC, NL
A. Sadeghi, Darmstadt, DE
I. Brown, Oxford, GB
F. Valgimigli, Metarini, IT
A. Guiseppi-Elie, Clemson, USA
S. Khayat, UFM, Iran
Q. Tan, Shanghai, China
Panel : How real and urgent are the security/privacy threats for IMDs? Which IMDs?
Springer Book underway, to appear early 2013
http://si.epfl.ch/SPIMD
Workshop on Security and Privacy in Implanted Medical Devices April 1, 2011 EPFL, Lausanne, Switzerland
Global cross-disciplinary efforts needed!

## Slide 23
SHARPS is a multi-institutional and multidisciplinary research project, supported by the Office of the National Coordinator for Health Information Technology, aimed at reducing security and privacy barriers to the effective use of health information technology. The project is organized around three major healthcare environments:
    Electronic Health Records (EHR)
    Health Information Exchange (HIE)
    Telemedicine (TEL)
A multidisciplinary team of computer security, medical, and social science experts is developing security and privacy policies and technology tools to support electronic use and exchange of health information.
UIUC, Stanford, Berkeley, Dartmouth, CMU, JHU, Vanderbilt, NYU, Harvard/BethIsrael, Northwestern, UWash, UMass
sharps.org
in the USA…

## Slide 24: Conclusions
Implantable Medical Devices have unique challenges in Security and Privacy
  Critical assets
  Resource constraints (power/energy, size)
  Hard to maintain
  Long lifetime
  Human factors
  Security/Safety tradeoffs
But solutions can leverage unique aspects of IMDs
  Proximity, in-body location
  Data-rates
  Threat models
Need to work with IMD designers and users
Much work to be done
  Cyber-physical and cyber-human systems
  Many exciting new IMDs
  Many possible new threats

## Slide 25: Backup/Q&A slides

## Slide 26: Threat taxonomy
D. Kotz, A threat taxonomy for mHealth privacy, NetHealth 2011

## Slide 27: Smart pills
Raisin, a digestible, ingestible microchip, can be put into medicines and food. Chip is activated and powered by stomach acids and can transmit to an external receiver from within the body! Useful for tracking existence and location of drugs, nutrients, etc.
Proteus Biomedical
Ingestible Raisin microchip
“…there’s more silicon in a banana…” - Proteus CTO

## Slide 28: Bio-sensors for hemorrhaging trauma victims
A. Guiseppe-Elie, C3B, Clemson University (USA)
Implantable biosensor for monitoring lactate and glucose levels. Funded by the US Department of Defense
Developing a temporary implantable dual sensing element biochip with wireless transmission capabilities.
Applications in mass triage scenarios such as battlefields and natural disaster sites provide a means for medical personnel to make life saving decisions.
Low-cost, short life-time, rapid deployment, life-saving
Future applications in diabetes care,
transplant organ health, and intensive care.

## Slide 29: Security and Privacy Design Issues
System Requirements
    Sensor/Actuator Functionality, Software updates
    Communications: Data-rate (>100kbps), Range/Channel (BAN)
    Protocol Design: Asymmetric channel, ( Active RFID)
Design Constraints
    Power (battery-powered, harvested, or remote-powered device)
    Size, Bio-compatibility, calibration
    Long life-time, little maintenance, reliability
Security Analysis
        Assets: Human health and well-being, personal and health data
    Threats: Device cloning and counterfeiting, Eavesdropping, Physical Layer Detection and Identification,
Security Primitives
    Public and private key crypto, block and stream ciphers, TRNG, PUF
    Secure radios, Distance-bounding protocols, etc.

## Slide 30: Personalized Therapies with multiple IMDs
The Development of new Implantable Medical Devices
is a key-factor for succeeding in Personalized therapy
1.Drug/marker
detection
2.Data Analysis
3.Therapy