## Slide 1: SoK: Security and Privacy in Implantable Medical Devices and Body Area Networks
Fall 2014
Presenter: Kun Sun, Ph.D.
Michael Rushanan*, Denis Foo Kune, Colleen M. Swanson, Aviel D. Rubin, s&p14
Thank Michael for sharing the slides.

## Slide 2: What is an Implantable Medical Device?
The FDA strictly defines a medical device
Device
  Embedded system that can sense and actuate
Implantable
  Surgically placed inside of a patient’s body
Medical
  Provides diagnosis and therapy for numerous health conditions

## Slide 3: IMD are not your typical PCs

## Slide 4: IMD are not your typical PCs

## Slide 5: IMD are not your typical PCs
There exists resource limitations
  The battery limits computation and is not rechargeable
There are safety and utility concerns
  The IMD must be beneficial to the patient and elevate patient safety above all else
  Security and privacy mechanisms must not adversely affect the patient or therapy
Lack of security mechanisms may have severe consequences
IMD’s provide safety-critical operation
  Must fail-open in the context of an emergency

## Slide 6: Research Questions
How do we provide security and privacy mechanisms that adequately consider safety and utility?
When do we use traditional security and privacy mechanisms or invent new protocols?
How do we formally evaluate security and privacy mechanisms?
Novel attack surfaces

## Slide 7: A Healthcare Story
Alice
Cardiac Carl
Nurse
Patient

## Slide 8: Cardiac Carl’s Condition
Atrial Fibrillation
Implantable Cardioverter Defibrillator
His ICD is safety-critical
Cardiac Carl
Atrial Fib.

## Slide 9: Alice and Carl’s Relationship
Nurse
Alice
visits
accesses ICD w/ programmer
receives private data
adjusts therapy
Where are the security and privacy mechanisms?
Cardiac
Carl

## Slide 10: Alice and Carl’s Relationship
Mallory
Hacker Elite

## Slide 11: Alice Mallory and Carl’s Relationship
Nurse
Alice
Cardiac
Carl
Mallory
wireless communication
  [Halperin, S&P, 08], [Li, HealthCom, 11]
eavesdrop
forge
modify
jam

## Slide 12: Attack Surfaces
Cardiac
Carl

## Slide 13: Red Herring?
Insulin pump: a self-contained controller
Most wireless communication is one direction
  Read data from IMD
  No write to IMD (like Dick Cheney’s)
  malware
Similar to Industrial Control System
  Can read data from sensors
  No write to control the actuators
  Stuxnet on Iran’s nuclear plant

## Slide 14: Security and Privacy Mechanisms
Security and Privacy mechanisms exist in standards
  Medical Implant Communication Services
  Wireless Medical Telemetry Service
These mechanisms are optional
Interoperability might take priority of security
  [Foo Kune, MedCOMM, 12]

## Slide 15
Biometrics and Physiological Values
Out-of-Band
Distance Bounding
Software/Malware
External Devices
Anomaly Detection
Future Work
Telemetry Interface

## Slide 16: Research Challenges
Access to Implantable Medical Devices
  Is much harder then getting other components
Reproducibility
  Limited analysis of attacks and defenses
  Do not use meat-based human tissue simulators
  Do use a calibrated saline solution at 1.8 g/L at 21 ◦C
    The complete design is described in the ANSI/AAMI PC69:2007 standard [92, Annex G]

## Slide 17: Security and Privacy Mechanisms
Biometric and Physiological Values
  Key generation and agreement
Electrocardiogram (ECG)
  Heart activity signal
Interpulse interval
  Time between heartbeats

## Slide 18: H2H Authentication Protocol
Nurse
Alice
  [Rostami, CCS, 13]
Cardiac
Carl
measure ECG α
measure ECG β
send ECG measurement β
send ECG measurement α
TLS without certs
Notes: Intended for emergency authentication with no pre-established relationship
IPI measurements are one time values and may be used to authenticate every session
When Carl decommits to Alice she checks if the measurements are approximately the same.
  If it is she decommits for Carl

## Slide 19: H2H Authentication Protocol
Adversarial Assumptions
  Active attacker with full network control
  The attacker cannot:
    Compromise the programmer
    Engage in a denial-of-service
    Remotely measure ECG to weaken authentication
  [Rostami, CCS, 13]
Notes: Strong adversary

## Slide 20: Physiological Values as an Entropy Source
How do ECG-based protocols work in practice?
  Age, Exertion, Noise
ECG-based protocols rely on an analysis of ideal data in an unrealistic setting
  Data sample is close to their ideal distribution
  Very accurate estimate of distribution characteristics
  Extract randomness using the estimate on the same data sample
Observability
  Using video processing techniques to extract ECG-signals
  [Rostami, S&P, 2013] [Chang, HealthTech, 2012]
  [Poh, Biomedical Engineering, 11]

## Slide 21
Future Work

## Slide 22: Trusted Sensor Interface
Current systems trust their analog sensor inputs
This assumption may not always hold
Forging signals using electromagnetic interference
  Inject cardiac waveform
[Foo Kune, S&P, 2013]

## Slide 23: Neurosecurity
Neurostimulators
  What are the new attack surfaces
  What are the implications of recording and transmitting brainwaves
Brain computer interfaces
Cognitive recognition could leak:
  Passwords, personal information
[Martinovic, USENIX, 2012], [Denning, Neurosurg Focus, 09]

## Slide 24: Questions?
IMDs are becoming more common
  Improving patient outcome
Research gaps exists
  Software
  Sensor Interface
Areas for future work include
  Physiological values as an Entropy Source
  Trusted Sensor Interface
  Neurosecurity