EKG Sensor System with Wireless Network Capabilities

Table of Contents

Theory of Operation of Analog Sensor Board……………………….……………………………1

Validation and Testing of Analog Sensor Board…………………………………………………..5

Simulation Results……………………………………………………………………...…5

Hardware Testing Flow Chart……………………...…………………………………….22

Simulation and Hardware Comparison with High Frequency Inputs……………………23

Hardware Results with Low Frequency Inputs…………………………………………..35

EKG Bill of Materials……………………………………………………………………………99

INA 321Datasheet………………………………………………………………………………100

OPA4336 Datasheet…………………………………………………………………………….122

Validation and Testing of Wireless Sensor Network……………………………………………131

Testing Flow Chart……………………………………………………………………...131

Single and Multi-Hop Wireless Network……………………………………………….132

Network Topology……………………………………………………………………...139

Sampling Rate………………………………………………………………………..…143

MICA2 Mote Battery Test……………………………………………………………...148

Signal Reconstruction…………………………………………………………………..149

MICA2 Datasheet………………………………………………………………………………163

MIB520 Datasheet……………………………………………………………………………...165

MDA100CB…………………………………………………………………………………….166

Integration Testing……………………………………………………………………………...170

Testing Flow Chart……………………………………………………………………...170

Full System Signal Reconstruction……………………………………………………..171

1 Hop Full System Test…………………………………………………………………185

Theory of Operation of Analog Sensor Board

Analog Sensor Schematic:

- Right Leg Drive Circuit

  - Purpose

    - The body may receive electromagnetic interference and act as an antenna. This interference is picked up by the lead attached at the lower chest and is driven back negatively into the body. By driving the interference back, the skin and the circuit may be set to a more common potential.

    - Overall, this section acts to loosely “ground” the patient and prevent harm while the device is in operation.

    - In general, a buffer amplifier's purpose is to keep one circuit from overloading another.

  - Configuration

    - Using a buffer amplifier, the signal received from the lower chest lead may be used as an input to the amplifier.  The second input is received by splitting the supply voltage in half across the voltage divider.

    - Combing these two signals in a negative feedback configuration and then feeding the output back to the body allows the circuit and skin to reach a common potential.

  - Expected Test Results

    - Ideally, a buffer amplifier's input voltage should equal it's output voltage due to the negative feedback configuration.

- INA321 Amplification

  - Purpose

    - The INA321 is the main amplifier in the circuit. Here, it amplifies the differential signal between the two inputs by a factor of 5. This factor of 5 is deemed as an adequate amplification of signals in the -5mV to 5mV range.

    - It's high Common Mode Rejection Ratio of 94dB also filters the common mode noise.

  - Configuration

    - The two leads attached to the patients chest feed their differential signal to the INA321 for amplification.

  - Expected Test Results

    - Amplification of the input signals by a factor of approximately 5.

- INA321 Support Amplifier

  - Purpose

    - A large amplification produces an output that is sensitive to variations in electrode to skin contact resistance. This is know as baseline wandering.  The OPA336 amplifier is used to eliminate this baseline wandering.

  - Configuration

    - The OPA336 amplifier integrates the DC content of the amplified signal and feeds it back to the INA321 through it's reference pin. This gives the INA321 a constant DC level at output regardless of skin contact resistance and eliminates baseline wandering.

- Filtering

  - Purpose

    - Both the third OPA336 from the left and the two filters at the end of the circuit behave as low pass filters. These filters pass low frequencies and weaken (cut-off) high frequencies.

    - In the final implementation, these filters will serve to eliminate additional noise from the EKG signal.

  - Configuration

    - The amplifier receives inputs from the INA321 and the voltage divider in the Right Leg Drive Circuit. The resulting output is an additional amplification of the signal from the INA321.

    - The cutoff frequency for the amplifier is found to be 33.898 Hz. Any frequency above this one in actual implementation should be cut off and not allowed to pass.

  - Expected Results

    - The output should reflect an amplification of the input signal from the INA321.

    - The output signal should not have any frequencies above 33.898Hz.

- Power Supply

  - Purpose

    - The INA321 amplifier operates with a supply voltage typically between 2.7V and 5.5V while the OPA4336 amplifier operates with a supply voltage typically between 2.3V and 5.5V. Coming from the Mica2, the analog sensor board will receive 3V DC voltage to power both amplifiers.

  - Configuration

    - By placing the capacitors in parallel with the voltage source, the capacitors remove any AC voltage that may enter the DC voltage as electrical interference.

  - Expected Results

    - Both amplifiers should operate given a voltage in their supply voltage range.

Validation and Testing of Analog Sensor Board

Simulation Results:

Note: All graphs show the results of transient analysis and graph voltage vs time.

- Test Unit 1: Right Leg Drive Circuit

    - Test 1: First verification of operation under normal operating conditions

      - Purpose:  Verify circuit operates as expected given input signals in the operational range and given input power supply in the range 2.3V to 5.5V.

      - Input Signals: 0.5V DC across a voltage divider

      - Input Power Supply: 3V DC

      - Conclusion:  Input voltages differential approximately equal output voltage as expected. The test was successful. Results are graphed below.

    - Test 2: Second verification of operation under normal operating conditions

      - Purpose:  Verify circuit operates as expected given input signals in the operational range and given input power supply in the range 2.3V to 5.5V.

      - Input Signals: 1V DC across a voltage divider

      - Input Power Supply: 2.8V DC

      - Conclusion:  Input voltages differential approximately equal output voltage as expected. The test was successful. Results are graphed below.

    - Test 3: Verification of non-operation when given voltage outside power supply requirements

      - Purpose: Verify the circuit will not operate as expected when a voltage supply is given that exceeds 7.5V

      - Input Signals: 0.5V DC across a voltage divider

      - Input Power Supply: 100V DC

      - Conclusion: Output voltage falls to zero; reflects the op-amps inability to operate. The test was successful. Results are graphed below.

    - Test 4: Verification of non-operation when given input voltage is outside typical input voltages

      - Purpose: Verify the circuit will not operate as expected when voltage across the signal input terminals exceeds 0.3V

      - Input Signals: 100V DC across a voltage divider

      - Input Power Supply: 3V DC

      - Conclusion: Input voltages not allowed to exceed approximately 3.75V; output voltage reflects differential of limited inputs.  The test was successful. Results are graphed below.

# Test Unit 2: INA321 Amplification

    - Test 1: Verification of correct operation of the INA321 precision amplifier under normal operating conditions

      - Purpose: Verify the circuit will operate as expected given input signals and an input voltage in its operating range

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm : 1mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Output voltage reflects an amplification of the two input signals by a factor of approximately 5. The test was successful. The results are graphed below.

    - Test 2: Verification of correct non-operation of the INA321 precision amplifier with an input voltage outside the normal range

      - Purpose: Verify the circuit will not operate correctly when it is supplied by a voltage source greater than 7.5V DC.

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm:  2mV AC with frequency of  2kHz

      - Input Power Supply: 100V DC

      - Conclusion: The given model did not accurately depict the INA321 being non-operational; the circuit still amplified the inputs signals as in Test 1. However, from other laboratory experience we know that a voltage source greater than what the op-amp is measured to handle will result in inoperation and damaged hardware. The results from the test are graphed below.

    - Test 3: Verification of  non-operation of the INA321 precision amplifier with an input signal outside the normal range

      - Purpose: Verify the circuit will not operate correctly when it is given an input voltage outside the range -0.5V to 0.5V.

      - Input Signals:

        - Right Arm: 100V AC with frequency of 2kHz

        - Left Arm: 2mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Output voltage amplifies signals up to its tolerance before cutting off, as expected. The results are graphed below.

# Test Unit 3: Low Pass Filtering

    - Test 1: Verification of correct operation of the low pass filter under normal operating conditions

      - Purpose: Verify the circuit will operate as expected given input signals and an input voltage in the op-amp's operating range.

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm : 1mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Output voltage reflects an amplification between two input signals. One input signal comes from the INA321 while the other comes from the voltage divider in the Right Leg Drive circuit. The approximate difference between the two input signals is the resulting output as seen below.

      - Note: This low pass filter is assumed operational if the first two stages of the circuit are operational

        - The INA321 precision amplifier verifies correct input signals

        - Correct operation of the first OPA336 used in the Right Leg Drive circuit verifies correct power supply voltage

# Test Unit 4: Final Filtering and Output

    - Test 1: Verification of correct operation of the final amplifier and filters under normal operating conditions

      - Purpose: Verify the circuit will operate as expected given input signals and an input voltage in the op-amp's operating range.

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm : 1mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Voltage obtained from the low pass filter remains as expected. The final output voltage through the final buffer amplifier is not as expected; the expected voltage should be similar to the voltage coming from the low pass filter. This test needs to be tested in the laboratory to observe if results are similar or different. Depending on the outcome, the buffer amplifier may need to be omitted in the final design.

      - Note: This section is assumed operational if the first two stages of the circuit are operational

        - The INA321 precision amplifier verifies correct input signals

        - Correct operation of the first OPA336 used in the Right Leg Drive circuit verifies correct power supply voltage

# Test Unit 5: Complete Circuit Testing

    - Test 1: Verification of correct operation given normal operation conditions

      - Purpose: Verify the circuit will operate as expected given input signals and an input voltage in the operating ranges of the op-amps

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm : 1mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Voltage obtained from the low pass filter remains as expected. The final output voltage through the final buffer amplifier is not as expected; the expected voltage should be similar to the voltage coming from the low pass filter. This test needs to be tested in the laboratory to observe if results are similar or different. Depending on the outcome, the buffer amplifier may need to be omitted in the final design.

    - Test 2: Verification of non-operation of the circuit given an input signal greater than it's normal operating range

      - Purpose: Verify the circuit will not operate correctly when the input voltages are not in the range -0.5V to 0.5V

      - Input Signals:

        - Right Arm: 100V AC with frequency of 2kHz

        - Left Arm:  2mV AC with frequency of  2kHz

      - Input Power Supply: 3V DC

      - Conclusion: Output voltage does not reflect any type of expected signal as seen below verifying the circuit is not operational.

    - Test 3: Verification of non-operation of the circuit given an input power supply greater than it's normal operating range

      - Purpose: Verify the circuit will not operate correctly when the input power supply is greater than 7.5V

      - Input Signals:

        - Right Arm: 5mV AC with frequency of 2kHz

        - Left Arm:  2mV AC with frequency of  2kHz

      - Input Power Supply: 100V DC

      - Conclusion: The final output of the circuit, shown below, does not match any expected output under normal operating conditions showing that the overall circuit was unable to operate correctly as expected.

Validation and Testing of Analog Sensor Board

Testing Flow Chart

Simulation and Laboratory Results Comparison

High Frequency Tests

Note: All graphs show the results of transient analysis and graph voltage vs. time

I. Global Testing Specifications

A. Purpose: To verify expected operation of the analog sensor at specific points and the final output with high frequency inputs; these results may be compared to the strictly simulation results in the previous section to give approximate expected values and behaviors. These tests also verify that the analog sensor board receives adequate power from the Mica2 mote.

B. Equipment:

(1) Two agilent function generators with one probe per generator

(2) One tektronix oscilloscope with three channels and one probe per channel

(3) One Analog Sensor PCB

(4) One Mica2 mote with attached DAQ

(5) Breadboard and assorted wires

    - Laboratory Setup:

Power from Mica2

Ground from Mica2

Test Point Probe to Oscilloscope

Function Generator Input Signal & Ground Connections

Input Signals to Board

Power to Board

Ground to Board

D. Input Signals:

(1) Left Arm: 1kHz square wave with 100mV amplitude

(2) Right Arm: 1kHz square wave with 500mV amplitude

(3) Right Leg: 1kHz square wave with 500mV amplitude

D. Input Power Supply: 3V DC (supplied by Mica2 mote)

F. Safety Precautions:

(1) Inputs (signal and power supply) were limited to the accepted operating 			ranges of the components to avoid damaging the hardware

(2) Power was cut to the analog sensor board during wiring modifications to 			  avoid shock

(3) Electrostatic resistant methods were followed to avoid damaging the 				hardware

(4) Power was cut to the analog sensor board during input signal changes to 			 avoid damaging the hardware

II. Test Unit 1: Right Leg Drive Circuit

  - Circuit Test Unit:

Schematic Unit

Corresponding parts on the PCB with Test Point 1 used for output

B. Simulation Results with Specified Inputs:

C. Laboratory Results with Specified Inputs:

Note: Orange represents the input from the right leg and blue represents the 				output from the right leg drive circuit

D. Conclusion: The measured results obtained from the hardware in laboratory reflect 		the results obtained from the simulator; the output is a mirror of the input due to the 		buffer amplifier configuration. The test was successful.

III. Test Unit 2: INA321 Amplification

A. Circuit Test Unit:

Schematic Unit

Corresponding parts on the PCB with Test Point 2 used for output

B. Simulation Results with Specified Inputs:

C. Laboratory Results with Specified Inputs:

D. Conclusion: Both the simulation and laboratory results depict an amplified version of 	the differential between the input signals. The results from laboratory are cleaner and 		avoid the spikes seen in the simulated results. This may be due to the additional 			precision and other specifications of the OPA336 amplifier used in the circuit versus the 	OPA335 amplifier used in the simulation. Overall the results obtained in the laboratory 		provide a better and more expected output than the simulation results resulting in a 		successful test for the hardware.

IV. Test Unit 3: Low Pass Filtering

A. Circuit Test Unit:

Schematic Unit

Corresponding parts on the PCB with Test Point 6 used for output

B. Simulation Results with Specified Inputs:

C. Laboratory Results with Specified Inputs:

D. Conclusion: Both the simulation and laboratory results depict a smaller amplified 		version of the differential coming from the INA321 as expected. Once again, the results 		from laboratory are cleaner and avoid the spikes seen in the simulated results. Overall, 		however, the results again approximately mirror the expected results marking the test as 		successful.

V: Test Unit 4: Final Filtering and Output

- Circuit Test Unit:

Schematic Unit

Corresponding parts on the PCB with PL6 used for output

B. Simulation Results with Specified Inputs:

C. Laboratory Results with Specified Inputs:

D. Conclusion: As with the previous section's test unit 4, the results obtained both from 		simulation and laboratory reflect undesirable values. This is due to the high frequency 		inputs, as will be shown in later tests. The low pass filters at the end of the circuit do not 	allow these higher frequencies through this part of the circuit resulting in the “garbage” 		data. If high frequency inputs are necessary during operation of the circuit, the final 		output must be taken from the previous point.

VI. Test Unit 5: Complete Circuit Testing

A. Circuit Test Unit:

Schematic Unit

B. Simulation Results with Specified Inputs:

C. Laboratory Results with Specified Inputs:

D. Conclusion: As with the previous section's test unit 4, the results obtained both from 	simulation and laboratory reflect undesirable values. This is due to the high frequency 		inputs, as will be shown in later tests. The low pass filters at the end of the circuit do not 	allow these higher frequencies through this part of the circuit resulting in the “garbage” 		data. If high frequency inputs are necessary during operation of the circuit, the final 		output must be taken from the previous point.

Laboratory Results

Low Frequency Tests: Square Wave Testing

I. Global Testing Specifications

A. Purpose: To verify expected operation of the analog sensor at specific points and the final output with low frequency inputs; these results may be compared to the strictly simulation results in the previous sections to give approximate expected values and behaviors. Low frequency tests were performed in preparation for integration testing because the Mica2 mote requires low frequency signals for adequate sampling and reconstruction.

B. Equipment:

(1) Two agilent function generators with one probe per generator

(2) One tektronix oscilloscope with three channels and one probe per channel

(3) One Analog Sensor PCB

(4) One Mica2 mote with attached DAQ

(5) Breadboard and assorted wires

C. Laboratory Setup:

Power from Mica2

Ground from Mica2

Test Point Probe to Oscilloscope

Function Generator Input Signal & Ground Connections

Input Signals to Board

Power to Board

Ground to Board

D.  Input Signals: The right leg, right arm, and left arm all received the same input signals. For these tests, this was a square wave with frequencies of 7.5Hz, 5Hz, 4Hz, 2Hz, 1Hz, and 0.5Hz to give a total of distinct tests.

E. Input Power Supply: 3V DC (supplied by Mica2 mote)

F. Safety Precautions:

(1) Inputs (signal and power supply) were limited to the accepted operating 			ranges of the components to avoid damaging the hardware

(2) Power was cut to the analog sensor board during wiring modifications to 			avoid shock

(3) Electrostatic resistant methods were followed to avoid damaging the 				hardware

(4) Power was cut to the analog sensor board during input signal changes to 			avoid damaging the hardware

II. 7.5Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test point. Here the output signal is reduced slightly after moving through the filter. The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output obtained is a reflection of the input signals as opposed to the previous tests with higher frequencies. Here, an amplified output signal measures 640mV proving that the circuit was successful in producing and filtering it's input signals as specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. Here, an amplified output 			signal measures 640mV proving that the circuit was successful in producing and 			filtering it's input signals as specified by the project specifications.

III. 5Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test point. Here the output signal is reduced slightly after moving through the filter. The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output obtained is a reflection of the input signals as opposed to the previous tests with higher frequencies. Here, an amplified output signal measures 680mV proving that the circuit was successful in producing and filtering it's input signals as specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as opposed to the previous tests with higher frequencies. An amplified output signal measures 680mV proving that the circuit was successful in producing and filtering it's input signals as specified by the project specifications.

III. 4Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output obtained is a reflection of the input signals as opposed to the previous tests with higher 	frequencies. Here, an amplified output signal measures 740mV proving that the circuit was successful in producing and filtering it's input signals as specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		measures 740mV proving that the circuit was successful in producing and filtering 		it's input signals as specified by the project specifications.

III. 2Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

Corresponding parts on the PCB with Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. 	Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful. The distortion is caused by the lower frequency 			value.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. 	Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful. The distortion is caused by the lower frequency 			value.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		proves that the circuit was successful in producing and filtering it's input signals 			as specified by the project specifications. The distortion in the signal is caused by 		the lower frequency signal.

III. 1Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

C. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output 			obtained is a reflection of the input signals as opposed to the previous tests with 			higher 	frequencies. Here, an amplified output signal measures 680mV proving 			that the circuit was successful in producing and filtering it's input signals as 			specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		proves that the circuit was successful in producing and filtering it's input signals 			as specified by the project specifications. The distortion in the signal is caused by 		the lower frequency signal.

III. 1Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Square wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Square wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

Corresponding parts on the PCB with Test Point 6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		proves that the circuit was successful in producing and filtering it's input signals 			as specified by the project specifications. The distortion in the signal is caused by 		the lower frequency signal.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Square wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		proves that the circuit was successful in producing and filtering it's input signals 			as specified by the project specifications. The distortion in the signal is caused by 		the lower frequency signal.

Laboratory Results

Low Frequency Tests: Sine Wave Testing

I. Global Testing Specifications

A. Purpose: To verify expected operation of the analog sensor at specific points and the final output with low frequency inputs; these results may be compared to the strictly simulation results in the previous sections to give approximate expected values and behaviors. Low frequency tests were performed in preparation for integration testing because the Mica2 mote requires low frequency signals for adequate sampling and reconstruction.

B. Equipment:

(1) Two agilent function generators with one probe per generator

(2) One tektronix oscilloscope with three channels and one probe per channel

(3) One Analog Sensor PCB

(4) One Mica2 mote with attached DAQ

(5) Breadboard and assorted wires

C. Laboratory Setup:

Power from Mica2

Ground from Mica2

Test Point Probe to Oscilloscope

Function Generator Input Signal & Ground Connections

Input Signals to Board

Power to Board

Ground to Board

D.  Input Signals: The right leg, right arm, and left arm all received the same input signals. For these tests, this was a sine wave with frequencies of 7.5Hz, 5Hz, 4Hz, 2Hz, 1Hz, and 0.5Hz to give a total of distinct tests.

E. Input Power Supply: 3V DC (supplied by Mica2 mote)

F. Safety Precautions:

(1) Inputs (signal and power supply) were limited to the accepted operating 			ranges of the components to avoid damaging the hardware

(2) Power was cut to the analog sensor board during wiring modifications to 			avoid shock

(3) Electrostatic resistant methods were followed to avoid damaging the 				 hardware

(4) Power was cut to the analog sensor board during input signal changes to 			avoid damaging the hardware

II. 7.5Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

Corresponding parts on the PCB with Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output 			obtained is a reflection of the input signals as opposed to the previous tests with 			higher 	frequencies. Here, an amplified output signal measures 580mV proving 			that the circuit was successful in producing and filtering it's input signals as 			specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 7.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. Here, an amplified output 			signal measures 580mV proving that the circuit was successful in producing and 			filtering it's input signals as specified by the project specifications.

III. 5Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output 			obtained is a reflection of the input signals as opposed to the previous tests with 			higher 	frequencies. Here, an amplified output signal measures 560mV proving 			that the circuit was successful in producing and filtering it's input signals as 			specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		measures 560mV proving that the circuit was successful in producing and filtering 		it's input signals as specified by the project specifications.

III. 4Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. 	Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: Due to the lower frequencies used in these test, the final output 			obtained is a reflection of the input signals as opposed to the previous tests with 			higher frequencies. Here, an amplified output signal measures 580mV proving 			that the circuit was successful in producing and filtering it's input signals as 			specified by the project specifications.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 4Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. An amplified output signal 		measures 580mV proving that the circuit was successful in producing and filtering 		it's input signals as specified by the project specifications.

III. 2Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful. The distortion is caused by the lower frequency 			value.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. 	Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful. The distortion is caused by the lower frequency 			value.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 2Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. This amplified signal 			proves that this test was successful.

III. 1Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B.  INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. This amplified signal 			proves that this test was successful.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 1Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. This amplified signal 			proves that this test was successful.

III. 1Hz Frequency Test

A. Right Leg Drive Circuit

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 1

(2) Input Signal: Sine wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: As expected, the output from the right leg drive circuit mirrors the input into the right leg; therefore, the test was successful.

B. INA321 Amplification

(1) Circuit Test Unit:

PCB with components under test outlined; output taken from Test Point 2

(2) Input Signal: Sine wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained at this test point depicts the amplified version of the differential signal between the two inputs. This is the expected behavior as has been shown in previous testing thereby ruling this test as a success.

C. Low Pass Filtering

(1) Circuit Test Unit:

PCB with components under test outlined; Test Point 6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The output obtained reflects the expected output from this test 			point. Here the output signal is reduced slightly after moving through the filter. 			The test was proven successful.

D. Final Filtering and Output

(1) Circuit Test Unit:

PCB with components under test outlined; PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. This amplified signal 			proves that this test was successful.

E. Complete Circuit Testing

(1) Circuit Test Unit:

PCB with PL6 used for output

(2) Input Signal: Sine wave with 500mV amplitude at 0.5Hz

(3) Laboratory Results with Specified Inputs:

(4) Conclusion: The final output obtained is a reflection of the input signals as 			opposed to the previous tests with higher frequencies. This amplified signal 			proves that this test was successful.

Bill of Materials: EKG Analog Board

| Quantity | Designator | Component | Package | Manufacturer Part # |
| --- | --- | --- | --- | --- |
| 1 | U2 | INA321 | 8-MSOP | INA321E/250 |
| 1 | U1 | OPA4336 | 16-SSOP | OPA4336EA/250 |
| 2 | R3, R4 | Resistor 100k, 5% | 0805 | CRCW0805100KJNEB |
| 3 | R7, R10, R11 | Resistor 1M, 5% | 0805 | CRCW08051M00JNEA |
| 3 | R1, R2, R12 | Resistor 47k, 5% | 0805 | CRCW080547K0JNEA |
| 2 | R5, R6 | Resistor 2.2M, 5% | 0805 | CRCW08052M20JNEA |
| 1 | R8 | Resistor 3.3k, 5% | 0805 | CRCW08053K30JNEA |
| 1 | R13 | Resistor 4.7k, 5% | 0805 | CRCW08054K70JNEA |
| 1 | R9 | Resistor 806, 5% | 0805 | CRCW0805806RFKEA |
| 2 | C5, C6 | Capacitor 0.1µF | 0805 | VJ0805Y104KXARW1BC |
| 2 | C1, C4 | Capacitor 0.1µF | 1210 | FCP1210C104G-G3 |
| 1 | C2 | Capacitor 4.7nF | 1206 | C1206C472K1RAC7025 |
| 1 | C3 | Capacitor 1µF | 1210 | ECP-U1C105MA5 |
| 3 | CONN1, CONN2, CONN3 | Connector | SMA/MCX | Not Bought at this Time |

These quantities are based on one sensor design.  All ordering of parts must be the number of sensors being designed times the quantities shown above.

Validation and Testing of Wireless Sensor Network

Testing Flow Chart

Single and Multi-Hop Wireless Sensor Network

Test Unit 1: 1-Hop Wireless Sensor Network

Purpose:

Test the sampling rate of MICA2 and MDA100CB acquisition board.

Test the ability to graphically represent the data transmitted and received.

Test the frequency band at which the data is being transmitted at.

Description:

The test done uses the MIB520 base station which is represented in the below figure as the Gateway (GW).  Attached to the gateway is one of the MICA2 motes used to receive the incoming data.  There are two other motes in this test represented as Nodes 1 and 2.  Node 1 during the whole test stays at room temperature.  Node 2 first starts out at room temperature, but then is placed in the freezer.  Upon stabilizing at 0ºC the node is then removed from the freezer and then stabilizes back to room temperature.  The figure below shows the two Nodes prior to Node 2 entering the freezer.

The below figure shows the topology of the network along with the temperature data of Node 2, when the node had stabilized at 0ºC.

The figure below shows the live data that was obtained for Node2 through the Moteview software provided by Crossbow.

The figure below shows the topology of the network and the resulting temperature data after Node2 had stabilized back around room temperature.

The Moteview Software also allows for the data to be exported to an excel file.  This allows for the data to be post-processed and analyzed.  Shown below is a small exert from the data obtained during the temperature test.

| Id | Time | voltage [V] | temp [C] | light |
| --- | --- | --- | --- | --- |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.92 | 923 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.92 | 923 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.92 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.92 | 924 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.838 | 925 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.838 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.838 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.838 | 924 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 925 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 925 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 925 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.838 | 924 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 925 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 925 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.757 | 924 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.838 | 884 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.838 | 908 |
| 2 | 10/9/2010 10:51 | 3.1948 | 20.92 | 915 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.92 | 833 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.757 | 757 |
| 2 | 10/9/2010 10:51 | 3.1866 | 20.431 | 567 |
| 2 | 10/9/2010 10:52 | 3.1866 | 20.186 | 262 |
| 2 | 10/9/2010 10:52 | 3.1866 | 19.941 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 19.778 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 19.452 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 19.289 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 19.044 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 18.717 | 0 |
| 2 | 10/9/2010 10:52 | 3.1866 | 18.391 | 0 |

Notice that the ID shows 2 which correctly identifies the node being tested.  Also 		notice the highlighted cells. At this time the light sensor shows no reading which 	represents the sensor going into the freezer.  Next notice the general trend of the 	temperature data after the first sign of darkness by the light sensor.

Conclusion:

The Moteview software allows for the network topology to be graphically 	represented.  Not only does the topology show the data, but a chart can be obtained in the 	software to show the live data.  The sampling rate was set by simply setting the data rate 	to 0.25 seconds.  The frequency band was tested by downloading the program onto the 	mote with the 916MHz frequency band set for communication.  If any device was set to a 	different frequency band the motes/nodes would not be able to effectively communicate.

Test Unit 2: Multi-Hop Wireless Sensor Network

Purpose:

Test the ability to achieve a multi-hop network

Test the distance between hops.

Description:

The test done uses the MIB520 base station which is represented in the below figure as the Gateway (GW).  Attached to the gateway is one of the MICA2 motes used to receive the incoming data.  There are two other motes in this test represented as Nodes 1 and 2.  To correctly show a multi-hop sensor network, Node 2 must use Node 1 to completely send its data to the gateway or base station.  The figure below shows the successful two hop network.

Conclusion:

In conclusion,  the graphical representation provided by the Moteview software shows that Node2 has to use Node1 to access the gateway.  The test also proved that inside Houser Hall a range of 20-40m was achieved before needing to a different mote to hop to the gateway.  Whenever the motes were outside the distance increased to a range of 50-100m all depending upon obstacles in the way.  The indoor range is drastically reduced compared to the outdoor range do to the walls, doors, and other obstacles that cause interference between the nodes.

Network Topology Test

The network topology test consisted of starting up a wireless sensor network with four different nodes.  All four of these nodes were initially connected through the base station. The figure below shows the initial setup.

After successfully setting up the network, several different topologies were desired to show full functionality and diversity of the network.  These different topologies can be seen below as four different cases.

Case 1:

The figure below shows the network topology of the network after setting up the network to resemble Case 1.  The topology was viewed through the Moteview Software.

Case 2:

The figure below shows the network topology of the network after setting up the network to resemble Case 2.  The topology was viewed through the Moteview Software.

Case 3:

The figure below shows the network topology of the network after setting up the network to resemble Case 3.  The topology was viewed through the Moteview Software.

Case 4:

The figure below shows the network topology of the network after setting up the network to resemble Case 4.  The topology was viewed through the Moteview Software.

Conclusion:

From the Network Topology Test it was determine that our network could handle the multiple nodes in many different configurations.  The Moteview software and code allowed for our network to configure in real time allowing for any configuration or topology desired.  The Network Topology Test successfully showed the ability of our wireless network to change in real time.

Sampling Rate Test

The sampling rate test consisted of taking as many readings as possible within a given time.  The time used for each test was 30 seconds.  In each test, the number of samples were then divided by the time to determine the frequency at which the MICA2 mote was sampling.  Four tests were completed and the results from the four test were then averaged to determine the sampling rate.

Test 1:

| Sample # | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 1 | 1.1576 | 1.3003 | 1.3511 | 1.3764 | 1.3955 |
| 2 | 1.1608 | 1.3067 | 1.3701 | 1.3796 | 1.3923 |
| 3 | 1.1386 | 1.2971 | 1.3542 | 1.3859 | 1.3986 |
| 4 | 1.1417 | 1.2908 | 1.3511 | 1.3733 | 1.3796 |
| 5 | 1.1608 | 1.2971 | 1.3542 | 1.3764 | 1.3891 |
| 6 | 1.1544 | 1.2908 | 1.3511 | 1.3701 | 1.3828 |
| 7 | 1.1703 | 1.3003 | 1.3542 | 1.3764 | 1.3955 |
| 8 | 1.1671 | 1.3067 | 1.3542 | 1.3859 | 1.3923 |
| 9 | 1.1639 | 1.294 | 1.3511 | 1.3733 | 1.3923 |
| 10 | 1.1513 | 1.294 | 1.3542 | 1.3764 | 1.3923 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 604 | 1.1481 | 1.2971 | 1.3606 | 1.3796 | 1.3955 |
| 605 | 1.1449 | 1.3003 | 1.3511 | 1.3859 | 1.3986 |
| 606 | 1.1513 | 1.2908 | 1.3606 | 1.3796 | 1.3859 |
| 607 | 1.1481 | 1.3067 | 1.3574 | 1.3796 | 1.3923 |
| 608 | 1.1544 | 1.2908 | 1.3606 | 1.3796 | 1.3859 |
| 609 | 1.1608 | 1.3003 | 1.3542 | 1.3733 | 1.3923 |
| 610 | 1.1386 | 1.2876 | 1.3479 | 1.3733 | 1.3828 |
| 611 | 1.1576 | 1.2971 | 1.3606 | 1.3859 | 1.3955 |
| 612 | 1.1513 | 1.3035 | 1.3606 | 1.3828 | 1.3923 |
| 613 | 1.1576 | 1.2971 | 1.3574 | 1.3764 | 1.3891 |
| 614 | 1.1449 | 1.3067 | 1.3574 | 1.3796 | 1.3923 |
| 615 | 1.1608 | 1.294 | 1.3542 | 1.3733 | 1.3891 |
| 616 | 1.2083 | 1.313 | 1.3701 | 1.3828 | 1.3955 |
| 617 | 1.1544 | 1.2971 | 1.3511 | 1.3701 | 1.3955 |
| 618 | 1.1386 | 1.2876 | 1.3542 | 1.3764 | 1.3923 |
| 619 | 1.1544 | 1.294 | 1.3511 | 1.3828 | 1.3923 |
| 620 | 1.1671 | 1.3035 | 1.3542 | 1.3764 | 1.3859 |
| 621 | 1.1639 | 1.3067 | 1.3669 | 1.3891 | 1.3986 |

Results:

Test 2:

| Sample # | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 1 | 1.1671 | 1.3098 | 1.332 | 1.3733 | 1.3891 |
| 2 | 1.1481 | 1.2908 | 1.3447 | 1.3669 | 1.3764 |
| 3 | 1.1449 | 1.3067 | 1.3606 | 1.3859 | 1.3606 |
| 4 | 1.1481 | 1.2971 | 1.3574 | 1.3859 | 1.3859 |
| 5 | 1.1576 | 1.294 | 1.3511 | 1.3733 | 1.3891 |
| 6 | 1.1354 | 1.2971 | 1.3542 | 1.3828 | 1.3891 |
| 7 | 1.1386 | 1.2971 | 1.3542 | 1.3828 | 1.4018 |
| 8 | 1.1576 | 1.294 | 1.3574 | 1.3796 | 1.3828 |
| 9 | 1.1449 | 1.294 | 1.3542 | 1.3796 | 1.3859 |
| 10 | 1.1513 | 1.3035 | 1.3637 | 1.3828 | 1.3511 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 594 | 1.1576 | 1.3035 | 1.3542 | 1.3796 | 1.3891 |
| 595 | 1.1576 | 1.2876 | 1.3574 | 1.3796 | 1.3955 |
| 596 | 1.1481 | 1.3035 | 1.3574 | 1.3828 | 1.3955 |
| 597 | 1.1735 | 1.3003 | 1.3606 | 1.3796 | 1.3955 |
| 598 | 1.2274 | 1.313 | 1.3606 | 1.3764 | 1.3891 |
| 599 | 1.1735 | 1.294 | 1.3542 | 1.3733 | 1.3859 |
| 600 | 1.1639 | 1.313 | 1.3606 | 1.3828 | 1.3859 |
| 601 | 1.1639 | 1.2971 | 1.3542 | 1.3796 | 1.3923 |
| 602 | 1.1671 | 1.294 | 1.3574 | 1.3796 | 1.3891 |
| 603 | 1.1449 | 1.3067 | 1.3574 | 1.3764 | 1.3891 |
| 604 | 1.1639 | 1.2971 | 1.3574 | 1.3733 | 1.3891 |
| 605 | 1.1576 | 1.2971 | 1.3511 | 1.3733 | 1.3923 |
| 606 | 1.1703 | 1.3067 | 1.3574 | 1.3796 | 1.3828 |
| 607 | 1.1576 | 1.3035 | 1.3637 | 1.3923 | 1.3923 |
| 608 | 1.1481 | 1.2971 | 1.3511 | 1.3796 | 1.3891 |
| 609 | 1.1449 | 1.3003 | 1.3574 | 1.3796 | 1.3891 |
| 610 | 1.1481 | 1.3035 | 1.3511 | 1.3733 | 1.3955 |
| 611 | 1.1608 | 1.3098 | 1.3637 | 1.3828 | 1.3955 |

Results:

Test 3:

| Sample # | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 1 | 1.1836 | 1.2898 | 1.3304 | 1.3491 | 1.3585 |
| 2 | 1.1805 | 1.2898 | 1.3429 | 1.3522 | 1.371 |
| 3 | 1.1805 | 1.2929 | 1.3366 | 1.3554 | 1.3679 |
| 4 | 1.1774 | 1.2991 | 1.3366 | 1.3585 | 1.371 |
| 5 | 1.168 | 1.296 | 1.3397 | 1.3554 | 1.3647 |
| 6 | 1.1866 | 1.2837 | 1.3369 | 1.3588 | 1.3713 |
| 7 | 1.1929 | 1.3025 | 1.3463 | 1.362 | 1.3713 |
| 8 | 1.1898 | 1.2929 | 1.3397 | 1.3522 | 1.3647 |
| 9 | 1.168 | 1.2898 | 1.3397 | 1.3585 | 1.3647 |
| 10 | 1.1866 | 1.2962 | 1.3432 | 1.3588 | 1.3807 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 503 | 1.1678 | 1.2993 | 1.34 | 1.362 | 1.3713 |
| 504 | 1.1929 | 1.3025 | 1.3526 | 1.3651 | 1.3807 |
| 505 | 1.168 | 1.2898 | 1.3366 | 1.3585 | 1.3741 |
| 506 | 1.1836 | 1.3054 | 1.3491 | 1.3679 | 1.3835 |
| 507 | 1.1804 | 1.2931 | 1.3463 | 1.362 | 1.3745 |
| 508 | 1.1835 | 1.2868 | 1.3369 | 1.362 | 1.3776 |
| 509 | 1.1711 | 1.2991 | 1.3522 | 1.3679 | 1.3679 |
| 510 | 1.1961 | 1.296 | 1.3397 | 1.3616 | 1.3835 |
| 511 | 1.1866 | 1.2993 | 1.3432 | 1.362 | 1.3776 |
| 512 | 1.1711 | 1.2773 | 1.3335 | 1.3585 | 1.371 |
| 513 | 1.1742 | 1.296 | 1.3335 | 1.3616 | 1.371 |
| 514 | 1.1929 | 1.2962 | 1.3463 | 1.3682 | 1.3839 |
| 515 | 1.1678 | 1.2899 | 1.3338 | 1.3526 | 1.3713 |
| 516 | 1.1772 | 1.3056 | 1.3494 | 1.362 | 1.3776 |
| 517 | 1.1867 | 1.2929 | 1.3335 | 1.3522 | 1.3647 |
| 518 | 1.1774 | 1.2929 | 1.3397 | 1.3647 | 1.371 |
| 519 | 1.1678 | 1.2899 | 1.3494 | 1.3651 | 1.3713 |
| 520 | 1.1866 | 1.2962 | 1.3463 | 1.3713 | 1.3807 |

Results:

Test 4:

| Sample # | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 1 | 1.1544 | 1.294 | 1.3637 | 1.3859 | 1.3891 |
| 2 | 1.1544 | 1.2971 | 1.3606 | 1.3828 | 1.3955 |
| 3 | 1.1481 | 1.3003 | 1.3542 | 1.3764 | 1.3859 |
| 4 | 1.1481 | 1.2908 | 1.3637 | 1.3828 | 1.3891 |
| 5 | 1.1513 | 1.2971 | 1.3574 | 1.3859 | 1.3955 |
| 6 | 1.1544 | 1.2971 | 1.3606 | 1.3828 | 1.3986 |
| 7 | 1.1388 | 1.2875 | 1.3571 | 1.3824 | 1.3824 |
| 8 | 1.1449 | 1.2971 | 1.3574 | 1.3796 | 1.3923 |
| 9 | 1.1322 | 1.2971 | 1.3542 | 1.3764 | 1.3923 |
| 10 | 1.1386 | 1.3035 | 1.3637 | 1.3859 | 1.3986 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 580 | 1.1291 | 1.3035 | 1.3669 | 1.3859 | 1.3955 |
| 581 | 1.1291 | 1.3035 | 1.3606 | 1.3796 | 1.3923 |
| 582 | 1.1513 | 1.2971 | 1.3637 | 1.3859 | 1.4018 |
| 583 | 1.1449 | 1.2971 | 1.3606 | 1.3764 | 1.3891 |
| 584 | 1.1322 | 1.3003 | 1.3606 | 1.3891 | 1.3955 |
| 585 | 1.1322 | 1.3003 | 1.3574 | 1.3859 | 1.3955 |
| 586 | 1.1354 | 1.3035 | 1.3669 | 1.3859 | 1.3986 |
| 587 | 1.1513 | 1.3003 | 1.3606 | 1.3828 | 1.3891 |
| 588 | 1.1386 | 1.3035 | 1.3542 | 1.3796 | 1.3955 |
| 589 | 1.1261 | 1.3001 | 1.3665 | 1.376 | 1.3855 |
| 590 | 1.1576 | 1.3067 | 1.3606 | 1.3796 | 1.3923 |
| 591 | 1.1513 | 1.3035 | 1.3606 | 1.3796 | 1.3859 |
| 592 | 1.1386 | 1.2971 | 1.3606 | 1.3828 | 1.3955 |
| 593 | 1.1259 | 1.3035 | 1.3542 | 1.3859 | 1.3923 |
| 594 | 1.1322 | 1.3035 | 1.3637 | 1.3796 | 1.3986 |
| 595 | 1.1417 | 1.3003 | 1.3637 | 1.3828 | 1.3923 |
| 596 | 1.1322 | 1.294 | 1.3574 | 1.3828 | 1.3891 |
| 597 | 1.1481 | 1.2971 | 1.3606 | 1.3859 | 1.3891 |

Results:

Conclusion:

It was concluded that our sampling rate was approximately 20Hz from our sampling rate test.  It was then determined that a sampling rate of 20Hz meant that we were sampling every 50ms as seen in the equation below.

MICA2 Mote Battery Test

The MICA2 Mote battery test consisted of creating a 1-hop wireless sensor network.  A MDA100CB data acquisition board was attached to the sensing MICA2.  The data from the MICA2 was then sent wirelessly to the base station.  The data was then collected and tabulated into a Moteview database.  The database was later exported to view the battery's voltage reading over 18 hours.  The test consisted of sampling light and temperature  every second to determine how quickly the battery drained.  The figure below shows the setup of the test.  The graph below shows the results from the MICA2 Mote Battery Test.

Conclusion:

The results from the MICA2 Mote Battery test shows that the MICA2 mote can continue sensing past 18 hours, which would be ideal to monitor a patient for a whole day.  From the graph and a continuation of this test it was determined that the battery's voltage level exponentially decreased and it would take longer to fully drain the battery from operating status.

Signal Reconstruction Test

Sync

Output

Waveform Generator

Base Station

RF Mote

USB

The signal reconstruction test consisted of applying a signal from a function generator directly to one of the ADC channels on the MDA100CB data acquisition board.  The data acquisition board was attached to the MICA2 mote which was directly connected to the MIB520 base station.  The figure below shows the setup of the test.  A sine wave and square wave were used at different frequencies to test the ability of the MICA2 mote and Moteview software to reconstruct the signal.  Frequencies of 0.5Hz, 1Hz, 2Hz, 4Hz, 5Hz and 7.5Hz were used to determine the limits of the software and mote.

Test 1: 0.5Hz Sine Wave

A 0.5Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 0.5Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1576 | 1.2908 | 1.3511 | 1.3764 | 0.89753 |
| 0.1 | 1.1671 | 1.2908 | 1.3415 | 1.3796 | 0.93559 |
| 0.15 | 1.1544 | 1.3035 | 1.3606 | 1.3859 | 0.97048 |
| 0.2 | 1.1671 | 1.3067 | 1.3511 | 1.3828 | 1.0244 |
| 0.25 | 1.1513 | 1.3067 | 1.3606 | 1.3955 | 1.0402 |
| 0.3 | 1.1513 | 1.3035 | 1.3574 | 1.3923 | 1.0751 |
| 0.35 | 1.1576 | 1.3067 | 1.3637 | 1.3923 | 1.0751 |
| 0.4 | 1.1513 | 1.3035 | 1.3606 | 1.3923 | 1.0878 |
| 0.45 | 1.1639 | 1.3003 | 1.3701 | 1.3986 | 1.0688 |
| 0.5 | 1.1544 | 1.3067 | 1.3574 | 1.3859 | 1.0624 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1386 | 1.2908 | 1.3511 | 1.3637 | 0.9229 |
| 9.65 | 1.1513 | 1.294 | 1.3447 | 1.3637 | 0.88485 |
| 9.7 | 1.1513 | 1.2876 | 1.3479 | 1.3637 | 0.84362 |
| 9.75 | 1.1671 | 1.2876 | 1.3352 | 1.3637 | 0.80239 |
| 9.8 | 1.1481 | 1.2971 | 1.3415 | 1.3606 | 0.75481 |
| 9.85 | 1.1449 | 1.294 | 1.3479 | 1.3733 | 0.69456 |
| 9.9 | 1.1449 | 1.2813 | 1.3352 | 1.3511 | 0.65333 |
| 9.95 | 1.1544 | 1.2908 | 1.3289 | 1.3542 | 0.61527 |
| 10 | 1.1735 | 1.3035 | 1.3415 | 1.3606 | 0.58038 |

Test 2: 0.5 Hz Square Wave

A 0.5Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 0.5Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1449 | 1.2876 | 1.3415 | 1.3637 | 0.48524 |
| 0.1 | 1.1513 | 1.2971 | 1.3415 | 1.3669 | 0.49158 |
| 0.15 | 1.1639 | 1.2908 | 1.3447 | 1.3733 | 0.48524 |
| 0.2 | 1.1576 | 1.3003 | 1.3415 | 1.3637 | 0.49158 |
| 0.25 | 1.1386 | 1.294 | 1.3511 | 1.3796 | 0.48841 |
| 0.3 | 1.1354 | 1.2813 | 1.3384 | 1.3637 | 0.49158 |
| 0.35 | 1.1322 | 1.2845 | 1.3447 | 1.3669 | 0.48841 |
| 0.4 | 1.1354 | 1.2781 | 1.3384 | 1.3637 | 0.49158 |
| 0.45 | 1.1417 | 1.2845 | 1.3479 | 1.3701 | 0.48841 |
| 0.5 | 1.1671 | 1.3035 | 1.3606 | 1.4145 | 1.0846 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1544 | 1.2908 | 1.3606 | 1.3891 | 1.091 |
| 9.65 | 1.1639 | 1.294 | 1.3542 | 1.3891 | 1.0878 |
| 9.7 | 1.1798 | 1.3098 | 1.3574 | 1.3923 | 1.0846 |
| 9.75 | 1.1765 | 1.3228 | 1.3673 | 1.3991 | 1.0811 |
| 9.8 | 1.1608 | 1.3067 | 1.3606 | 1.3955 | 1.0846 |
| 9.85 | 1.1735 | 1.3162 | 1.3733 | 1.4018 | 1.0783 |
| 9.9 | 1.1735 | 1.3067 | 1.3701 | 1.3923 | 1.0846 |
| 9.95 | 1.1639 | 1.3098 | 1.3701 | 1.3923 | 1.0783 |
| 10 | 1.1608 | 1.3035 | 1.3637 | 1.3859 | 1.0846 |

Test 3: 1.0Hz Sine Wave

A 1.0 Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 1.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1671 | 1.3003 | 1.3574 | 1.3891 | 0.8785 |
| 0.1 | 1.183 | 1.3035 | 1.3606 | 1.3891 | 0.97048 |
| 0.15 | 1.1766 | 1.3162 | 1.3701 | 1.3986 | 1.0466 |
| 0.2 | 1.1925 | 1.3193 | 1.3733 | 1.4081 | 1.0783 |
| 0.25 | 1.1925 | 1.313 | 1.3701 | 1.3955 | 1.0783 |
| 0.3 | 1.183 | 1.313 | 1.3637 | 1.3859 | 1.0307 |
| 0.35 | 1.1861 | 1.3193 | 1.3733 | 1.3891 | 0.96096 |
| 0.4 | 1.1671 | 1.3098 | 1.3637 | 1.3796 | 0.85313 |
| 0.45 | 1.1671 | 1.3035 | 1.3542 | 1.3669 | 0.75164 |
| 0.5 | 1.1608 | 1.3035 | 1.3511 | 1.3606 | 0.64381 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1544 | 1.2971 | 1.3447 | 1.3701 | 0.58038 |
| 9.65 | 1.1671 | 1.294 | 1.3415 | 1.3733 | 0.6787 |
| 9.7 | 1.1639 | 1.2971 | 1.3415 | 1.3764 | 0.78019 |
| 9.75 | 1.1671 | 1.3003 | 1.3542 | 1.3764 | 0.89436 |
| 9.8 | 1.1544 | 1.2971 | 1.3574 | 1.3891 | 0.97999 |
| 9.85 | 1.1861 | 1.3098 | 1.3606 | 1.3955 | 1.0529 |
| 9.9 | 1.183 | 1.3162 | 1.3701 | 1.4018 | 1.0783 |
| 9.95 | 1.1798 | 1.313 | 1.3669 | 1.3955 | 1.0751 |
| 10 | 1.1639 | 1.3098 | 1.3669 | 1.3891 | 1.0244 |

Test 4: 1.0Hz Square Wave

A 1.0 Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 1.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.0688 | 1.2654 | 1.3415 | 1.3859 | 1.0846 |
| 0.1 | 1.0656 | 1.2623 | 1.3415 | 1.3828 | 1.0846 |
| 0.15 | 1.0529 | 1.2686 | 1.3384 | 1.3669 | 1.0846 |
| 0.2 | 1.0656 | 1.2718 | 1.3479 | 1.3701 | 1.0846 |
| 0.25 | 1.0593 | 1.2718 | 1.3384 | 1.3701 | 1.0815 |
| 0.3 | 1.0656 | 1.2718 | 1.3447 | 1.3733 | 1.0846 |
| 0.35 | 1.0656 | 1.2686 | 1.3447 | 1.3764 | 1.0846 |
| 0.4 | 1.072 | 1.2749 | 1.3511 | 1.3828 | 1.0846 |
| 0.45 | 1.0783 | 1.2749 | 1.3606 | 1.3891 | 1.0846 |
| 0.5 | 1.0498 | 1.2686 | 1.3511 | 1.3796 | 1.0846 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.0815 | 1.2813 | 1.3574 | 1.3828 | 1.0846 |
| 9.65 | 1.0751 | 1.2654 | 1.3447 | 1.3479 | 0.48524 |
| 9.7 | 1.0624 | 1.2527 | 1.3289 | 1.3511 | 0.48524 |
| 9.75 | 1.0561 | 1.2527 | 1.3257 | 1.3511 | 0.48524 |
| 9.8 | 1.0434 | 1.2527 | 1.332 | 1.3479 | 0.48524 |
| 9.85 | 1.0656 | 1.2559 | 1.3225 | 1.3511 | 0.48524 |
| 9.9 | 1.0624 | 1.2559 | 1.332 | 1.3574 | 0.48524 |
| 9.95 | 1.0529 | 1.2559 | 1.3257 | 1.3542 | 0.48524 |
| 10 | 1.0593 | 1.2591 | 1.3225 | 1.3511 | 0.48524 |

Test 5: 2.0 Hz Sine Wave

A 2.0 Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 2.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.183 | 1.3035 | 1.3511 | 1.3828 | 1.0846 |
| 0.1 | 1.1608 | 1.3067 | 1.3542 | 1.3733 | 1.0054 |
| 0.15 | 1.1449 | 1.2876 | 1.3415 | 1.3637 | 0.84996 |
| 0.2 | 1.1671 | 1.2845 | 1.3447 | 1.3606 | 0.67236 |
| 0.25 | 1.1671 | 1.294 | 1.3384 | 1.3606 | 0.53281 |
| 0.3 | 1.1671 | 1.2845 | 1.3415 | 1.3542 | 0.49158 |
| 0.35 | 1.1544 | 1.2876 | 1.3415 | 1.3574 | 0.55501 |
| 0.4 | 1.1513 | 1.294 | 1.3542 | 1.3764 | 0.70407 |
| 0.45 | 1.1449 | 1.3003 | 1.3574 | 1.3828 | 0.9229 |
| 0.5 | 1.1544 | 1.3003 | 1.3574 | 1.3955 | 1.0593 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1703 | 1.294 | 1.3542 | 1.3828 | 0.87533 |
| 4.65 | 1.1481 | 1.2908 | 1.3511 | 1.3637 | 0.66601 |
| 4.7 | 1.1513 | 1.2876 | 1.3384 | 1.3542 | 0.51061 |
| 4.75 | 1.1639 | 1.2876 | 1.3479 | 1.3637 | 0.50744 |
| 4.8 | 1.1544 | 1.2908 | 1.3479 | 1.3669 | 0.63747 |
| 4.85 | 1.1703 | 1.294 | 1.3447 | 1.3764 | 0.85313 |
| 4.9 | 1.1798 | 1.3067 | 1.3574 | 1.3923 | 1.0244 |
| 4.95 | 1.1671 | 1.3098 | 1.3606 | 1.3986 | 1.0846 |
| 5 | 1.183 | 1.3098 | 1.3669 | 1.3859 | 0.97682 |

Test 6: 2.0 Hz Square Wave

A 2.0 Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 2.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1576 | 1.2876 | 1.3352 | 1.3542 | 0.49158 |
| 0.1 | 1.1449 | 1.2908 | 1.3384 | 1.3574 | 0.49158 |
| 0.15 | 1.1449 | 1.3003 | 1.3637 | 1.4081 | 1.0846 |
| 0.2 | 1.1608 | 1.3098 | 1.3606 | 1.4018 | 1.0783 |
| 0.25 | 1.1513 | 1.3098 | 1.3637 | 1.405 | 1.0846 |
| 0.3 | 1.1513 | 1.3098 | 1.3574 | 1.3891 | 1.0783 |
| 0.35 | 1.1671 | 1.3003 | 1.3542 | 1.3542 | 0.49158 |
| 0.4 | 1.1354 | 1.2845 | 1.3447 | 1.3542 | 0.48524 |
| 0.45 | 1.1513 | 1.2876 | 1.3479 | 1.3574 | 0.49158 |
| 0.5 | 1.1513 | 1.3003 | 1.3479 | 1.3606 | 0.48524 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1671 | 1.3035 | 1.3542 | 1.3986 | 1.0815 |
| 4.65 | 1.1513 | 1.3098 | 1.3606 | 1.3923 | 1.0846 |
| 4.7 | 1.1703 | 1.3193 | 1.3733 | 1.4081 | 1.0815 |
| 4.75 | 1.1576 | 1.313 | 1.3606 | 1.3923 | 1.0846 |
| 4.8 | 1.1513 | 1.3035 | 1.3606 | 1.3637 | 0.48841 |
| 4.85 | 1.1576 | 1.2876 | 1.3479 | 1.3542 | 0.48841 |
| 4.9 | 1.1513 | 1.2876 | 1.3511 | 1.3701 | 0.49158 |
| 4.95 | 1.1544 | 1.2908 | 1.332 | 1.3542 | 0.48841 |
| 5 | 1.1386 | 1.2876 | 1.3447 | 1.3637 | 0.49158 |

Test 7: 4.0 Hz Sine Wave

A 4.0 Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 4.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1544 | 1.3035 | 1.3479 | 1.3606 | 0.50744 |
| 0.1 | 1.1417 | 1.2971 | 1.3606 | 1.3828 | 0.87216 |
| 0.15 | 1.1703 | 1.3067 | 1.3669 | 1.3955 | 1.072 |
| 0.2 | 1.1608 | 1.294 | 1.3606 | 1.3796 | 0.76433 |
| 0.25 | 1.1513 | 1.2971 | 1.3447 | 1.3606 | 0.48524 |
| 0.3 | 1.1513 | 1.2971 | 1.3542 | 1.3701 | 0.75799 |
| 0.35 | 1.1513 | 1.3098 | 1.3669 | 1.3986 | 1.072 |
| 0.4 | 1.1449 | 1.3003 | 1.3606 | 1.3859 | 0.88167 |
| 0.45 | 1.1544 | 1.2908 | 1.3606 | 1.3733 | 0.50744 |
| 0.5 | 1.1291 | 1.2876 | 1.3447 | 1.3669 | 0.65016 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1639 | 1.3003 | 1.3637 | 1.3891 | 1.0307 |
| 2.15 | 1.1671 | 1.3035 | 1.3542 | 1.3701 | 0.65016 |
| 2.2 | 1.1576 | 1.2908 | 1.3511 | 1.3542 | 0.50744 |
| 2.25 | 1.1576 | 1.2908 | 1.3479 | 1.3796 | 0.87216 |
| 2.3 | 1.1481 | 1.2971 | 1.3574 | 1.3891 | 1.0751 |
| 2.35 | 1.1608 | 1.2876 | 1.3542 | 1.3701 | 0.76116 |
| 2.4 | 1.1735 | 1.2845 | 1.3447 | 1.3542 | 0.5011 |
| 2.45 | 1.1608 | 1.2781 | 1.3447 | 1.3606 | 0.62161 |
| 2.5 | 1.1671 | 1.294 | 1.3447 | 1.3764 | 0.96096 |

Test 8: 4.0 Hz Square Wave

A 4.0 Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 4.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1766 | 1.3067 | 1.3637 | 1.3986 | 1.0878 |
| 0.1 | 1.1671 | 1.3003 | 1.3511 | 1.3542 | 0.48524 |
| 0.15 | 1.1386 | 1.2845 | 1.3352 | 1.3511 | 0.49158 |
| 0.2 | 1.1576 | 1.2971 | 1.3637 | 1.405 | 1.0783 |
| 0.25 | 1.1671 | 1.3003 | 1.3637 | 1.405 | 1.0846 |
| 0.3 | 1.1703 | 1.2971 | 1.3511 | 1.3511 | 0.48841 |
| 0.35 | 1.1449 | 1.2971 | 1.3447 | 1.3606 | 0.49158 |
| 0.4 | 1.1735 | 1.3162 | 1.3701 | 1.405 | 1.0815 |
| 0.45 | 1.1544 | 1.3162 | 1.3701 | 1.4113 | 1.0846 |
| 0.5 | 1.1481 | 1.294 | 1.3542 | 1.3733 | 0.48841 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1608 | 1.294 | 1.3415 | 1.3542 | 0.48841 |
| 2.15 | 1.1449 | 1.3003 | 1.3574 | 1.3923 | 1.0846 |
| 2.2 | 1.1608 | 1.313 | 1.3701 | 1.4018 | 1.0783 |
| 2.25 | 1.1544 | 1.294 | 1.3447 | 1.3606 | 0.49158 |
| 2.3 | 1.1449 | 1.2845 | 1.3384 | 1.3479 | 0.49158 |
| 2.35 | 1.1354 | 1.2876 | 1.3447 | 1.3542 | 0.48841 |
| 2.4 | 1.1671 | 1.294 | 1.3511 | 1.3986 | 1.0815 |
| 2.45 | 1.1671 | 1.3098 | 1.3574 | 1.4018 | 1.0783 |
| 2.5 | 1.1735 | 1.3098 | 1.3637 | 1.3955 | 1.0751 |

Test 9: 5.0 Hz Sine Wave

A 5.0 Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 5.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1576 | 1.3003 | 1.3637 | 1.3955 | 1.018 |
| 0.1 | 1.1608 | 1.2971 | 1.3542 | 1.3606 | 0.4789 |
| 0.15 | 1.1576 | 1.3035 | 1.3479 | 1.3701 | 0.74847 |
| 0.2 | 1.1639 | 1.3035 | 1.3637 | 1.3986 | 1.0402 |
| 0.25 | 1.1386 | 1.294 | 1.3511 | 1.3669 | 0.70724 |
| 0.3 | 1.1576 | 1.294 | 1.3511 | 1.3606 | 0.50744 |
| 0.35 | 1.1449 | 1.3035 | 1.3542 | 1.3923 | 1.0307 |
| 0.4 | 1.1639 | 1.3035 | 1.3574 | 1.3923 | 0.8785 |
| 0.45 | 1.1417 | 1.294 | 1.3479 | 1.3511 | 0.51378 |
| 0.5 | 1.1449 | 1.2749 | 1.3415 | 1.3637 | 0.79922 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1417 | 1.2908 | 1.3384 | 1.3511 | 0.55184 |
| 2.15 | 1.1417 | 1.2971 | 1.3511 | 1.3891 | 0.94193 |
| 2.2 | 1.1703 | 1.3067 | 1.3542 | 1.3891 | 0.97999 |
| 2.25 | 1.1481 | 1.294 | 1.3479 | 1.3606 | 0.49792 |
| 2.3 | 1.1354 | 1.2845 | 1.3447 | 1.3669 | 0.75164 |
| 2.35 | 1.1513 | 1.2971 | 1.3542 | 1.3986 | 1.072 |
| 2.4 | 1.1671 | 1.3003 | 1.3479 | 1.3733 | 0.73579 |
| 2.45 | 1.1513 | 1.2876 | 1.3415 | 1.3542 | 0.45035 |
| 2.5 | 1.1417 | 1.2845 | 1.3415 | 1.3701 | 0.75799 |

Test 10: 5.0 Hz Square Wave

A 5.0 Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 5.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1703 | 1.3098 | 1.3701 | 1.4113 | 1.0846 |
| 0.1 | 1.1576 | 1.3003 | 1.3479 | 1.3542 | 0.48841 |
| 0.15 | 1.183 | 1.3003 | 1.3479 | 1.3542 | 0.49158 |
| 0.2 | 1.1703 | 1.3162 | 1.3637 | 1.4081 | 1.0815 |
| 0.25 | 1.1861 | 1.3067 | 1.3574 | 1.3574 | 0.48841 |
| 0.3 | 1.1798 | 1.3098 | 1.3511 | 1.3542 | 0.49158 |
| 0.35 | 1.1925 | 1.3257 | 1.3637 | 1.4081 | 1.0815 |
| 0.4 | 1.1861 | 1.3193 | 1.3701 | 1.405 | 1.0846 |
| 0.45 | 1.1766 | 1.3067 | 1.3479 | 1.3511 | 0.48841 |
| 0.5 | 1.1671 | 1.3003 | 1.3574 | 1.3955 | 1.0846 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1735 | 1.313 | 1.3701 | 1.4113 | 1.0783 |
| 2.15 | 1.1671 | 1.2971 | 1.3511 | 1.3574 | 0.49158 |
| 2.2 | 1.1639 | 1.3003 | 1.3511 | 1.3542 | 0.48841 |
| 2.25 | 1.1893 | 1.3162 | 1.3733 | 1.4177 | 1.0846 |
| 2.3 | 1.1925 | 1.3225 | 1.3733 | 1.4145 | 1.0815 |
| 2.35 | 1.1671 | 1.3035 | 1.3542 | 1.3511 | 0.49158 |
| 2.4 | 1.1639 | 1.3098 | 1.3669 | 1.3955 | 1.0815 |
| 2.45 | 1.1703 | 1.313 | 1.3606 | 1.4018 | 1.0846 |
| 2.5 | 1.1735 | 1.3067 | 1.3574 | 1.3574 | 0.48841 |

Test 11: 7.5 Hz Sine Wave

A 7.5 Hz Sine Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 7.5 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1449 | 1.294 | 1.3479 | 1.3511 | 0.48841 |
| 0.1 | 1.1449 | 1.294 | 1.3542 | 1.3923 | 1.0212 |
| 0.15 | 1.1544 | 1.2876 | 1.3479 | 1.3606 | 0.63113 |
| 0.2 | 1.1766 | 1.2971 | 1.3352 | 1.3606 | 0.69139 |
| 0.25 | 1.1639 | 1.3003 | 1.3415 | 1.3796 | 0.97048 |
| 0.3 | 1.1576 | 1.2813 | 1.3384 | 1.3415 | 0.52012 |
| 0.35 | 1.1513 | 1.2971 | 1.3574 | 1.4018 | 1.0434 |
| 0.4 | 1.1386 | 1.2908 | 1.3415 | 1.3574 | 0.56453 |
| 0.45 | 1.1513 | 1.3003 | 1.3606 | 1.3828 | 0.91339 |
| 0.5 | 1.1417 | 1.2876 | 1.3511 | 1.3796 | 0.75481 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1703 | 1.3067 | 1.3606 | 1.3986 | 1.0371 |
| 2.15 | 1.1481 | 1.2876 | 1.3352 | 1.3542 | 0.51061 |
| 2.2 | 1.1449 | 1.2971 | 1.3542 | 1.3891 | 0.97999 |
| 2.25 | 1.1481 | 1.3003 | 1.3542 | 1.3637 | 0.63747 |
| 2.3 | 1.1576 | 1.3003 | 1.3542 | 1.3796 | 0.80873 |
| 2.35 | 1.1513 | 1.2971 | 1.3447 | 1.3701 | 0.84362 |
| 2.4 | 1.1576 | 1.2845 | 1.3384 | 1.3479 | 0.6121 |
| 2.45 | 1.1417 | 1.2908 | 1.3447 | 1.3891 | 1.1005 |
| 2.5 | 1.1608 | 1.2908 | 1.3479 | 1.3606 | 0.64064 |

Test 12: 7.5 Hz Square Wave

A 7.5 Hz Square Wave was applied from the function generator with 300mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 7.5 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1608 | 1.294 | 1.3511 | 1.3955 | 1.0149 |
| 0.1 | 1.1449 | 1.2908 | 1.3511 | 1.3637 | 0.39644 |
| 0.15 | 1.1386 | 1.2876 | 1.3415 | 1.3511 | 0.56453 |
| 0.2 | 1.1544 | 1.2971 | 1.3574 | 1.405 | 1.0212 |
| 0.25 | 1.1386 | 1.2845 | 1.3415 | 1.3511 | 0.5233 |
| 0.3 | 1.1576 | 1.3003 | 1.3511 | 1.3923 | 1.0244 |
| 0.35 | 1.1386 | 1.2908 | 1.3415 | 1.3415 | 0.41547 |
| 0.4 | 1.1544 | 1.294 | 1.3606 | 1.3796 | 1.0466 |
| 0.45 | 1.1735 | 1.3098 | 1.3574 | 1.3955 | 1.0593 |
| 0.5 | 1.1513 | 1.294 | 1.3415 | 1.3511 | 0.50427 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1671 | 1.3098 | 1.3637 | 1.4018 | 1.0942 |
| 2.15 | 1.1354 | 1.2845 | 1.3384 | 1.3447 | 0.45669 |
| 2.2 | 1.1481 | 1.3003 | 1.3669 | 1.4018 | 1.0973 |
| 2.25 | 1.1481 | 1.2749 | 1.3447 | 1.3542 | 0.45035 |
| 2.3 | 1.1576 | 1.294 | 1.3574 | 1.3955 | 1.0688 |
| 2.35 | 1.1608 | 1.2971 | 1.3606 | 1.3701 | 0.49475 |
| 2.4 | 1.1513 | 1.2908 | 1.3574 | 1.3859 | 1.0561 |
| 2.45 | 1.1576 | 1.3067 | 1.3701 | 1.4113 | 1.1037 |
| 2.5 | 1.1544 | 1.2971 | 1.3479 | 1.3542 | 0.45035 |

Conclusion

After performing signal reconstruction tests for the frequencies of 0.5 Hz, 1.0 Hz, 2.0 Hz, 4.0 Hz, 5.0 Hz, and 7.5Hz, the maximum signal frequency for the system was determined.  The maximum achievable sampling rate of 20.0 Hz limits the maximum measurable signal frequency to 4.0 Hz, for accurate signal reconstruction.  As shown in the test data, signals exceeding 4.0 Hz were not able to be accurately reconstructed.

Integration Testing

Full System Signal Reconstruction Test

EKG Sensor

Base Station

RF Mote

USB

Sync

Output

Waveform Generator

Full system reconstruction tests were performed by applying a signal from a function generation to the EKG sensor board.  The output signal was then linked to ADC6 of the MDA100CB data acquisition board.  The data acquisition board was attached to the MICA2 mote which was directly connected to the MIB520 base station.  The figure below shows the setup of the test. The MICA2 Mote supplies the power for the whole system.  Sine waves and square waves were used at different frequencies to test the ability of the MICA2 mote and Moteview software to reconstruct the signal through the EKG sensor board.  Frequencies of 0.5Hz, 1Hz, 2Hz, 4Hz, 5Hz and 7.5Hz were used to determine the system capabilities.

Test 1: 0.5Hz Sine Wave

A 0.5Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 0.5Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1988 | 1.3067 | 1.3574 | 1.3733 | 1.3003 |
| 0.1 | 1.2083 | 1.3162 | 1.3574 | 1.3796 | 1.2876 |
| 0.15 | 1.2083 | 1.313 | 1.3606 | 1.3733 | 1.2749 |
| 0.2 | 1.2083 | 1.3162 | 1.3637 | 1.3733 | 1.2591 |
| 0.25 | 1.1893 | 1.313 | 1.3606 | 1.3733 | 1.2464 |
| 0.3 | 1.202 | 1.3162 | 1.3606 | 1.3828 | 1.2369 |
| 0.35 | 1.1893 | 1.313 | 1.3606 | 1.3828 | 1.2242 |
| 0.4 | 1.2147 | 1.3193 | 1.3574 | 1.3796 | 1.2179 |
| 0.45 | 1.2083 | 1.3162 | 1.3606 | 1.3733 | 1.2083 |
| 0.5 | 1.1957 | 1.3193 | 1.3669 | 1.3828 | 1.2052 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1798 | 1.3003 | 1.3542 | 1.3764 | 1.3162 |
| 9.65 | 1.1798 | 1.3067 | 1.3511 | 1.3764 | 1.3193 |
| 9.7 | 1.1957 | 1.3067 | 1.3574 | 1.3764 | 1.3225 |
| 9.75 | 1.1893 | 1.3067 | 1.3574 | 1.3764 | 1.3225 |
| 9.8 | 1.1798 | 1.3067 | 1.3511 | 1.3701 | 1.3193 |
| 9.85 | 1.1735 | 1.3067 | 1.3542 | 1.3764 | 1.313 |
| 9.9 | 1.1957 | 1.3098 | 1.3574 | 1.3796 | 1.3035 |
| 9.95 | 1.1766 | 1.3098 | 1.3637 | 1.3764 | 1.294 |
| 10 | 1.1798 | 1.3098 | 1.3574 | 1.3733 | 1.2813 |

Test 2: 0.5 Hz Square Wave

A 0.5Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 0.5Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1957 | 1.3257 | 1.3859 | 1.4018 | 1.1481 |
| 0.1 | 1.1861 | 1.3225 | 1.3764 | 1.3923 | 1.1481 |
| 0.15 | 1.1703 | 1.313 | 1.3701 | 1.3891 | 1.1481 |
| 0.2 | 1.1798 | 1.3193 | 1.3733 | 1.3986 | 1.1481 |
| 0.25 | 1.1703 | 1.3193 | 1.3701 | 1.3891 | 1.1481 |
| 0.3 | 1.1671 | 1.313 | 1.3606 | 1.3859 | 1.1481 |
| 0.35 | 1.1703 | 1.313 | 1.3669 | 1.3891 | 1.1481 |
| 0.4 | 1.1639 | 1.3162 | 1.3701 | 1.3955 | 1.1481 |
| 0.45 | 1.183 | 1.3225 | 1.3637 | 1.3923 | 1.1481 |
| 0.5 | 1.1798 | 1.3035 | 1.3637 | 1.3891 | 1.1481 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1736 | 1.3001 | 1.3507 | 1.3634 | 0.9743 |
| 9.65 | 1.1703 | 1.3067 | 1.3542 | 1.3764 | 0.97999 |
| 9.7 | 1.1798 | 1.3003 | 1.3479 | 1.3733 | 0.97682 |
| 9.75 | 1.1671 | 1.3067 | 1.3574 | 1.3669 | 0.97999 |
| 9.8 | 1.1544 | 1.2971 | 1.3542 | 1.3701 | 0.97682 |
| 9.85 | 1.1766 | 1.313 | 1.3606 | 1.3733 | 0.97999 |
| 9.9 | 1.1576 | 1.3067 | 1.3669 | 1.3955 | 1.1639 |
| 9.95 | 1.183 | 1.3257 | 1.3796 | 1.4081 | 1.1639 |
| 10 | 1.1735 | 1.3193 | 1.3764 | 1.405 | 1.1639 |

Test 3: 1.0 Hz Sine Wave

A 1.0 Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 1.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.2115 | 1.3098 | 1.3511 | 1.3669 | 0.73579 |
| 0.1 | 1.2083 | 1.3098 | 1.3479 | 1.3669 | 0.75481 |
| 0.15 | 1.221 | 1.3098 | 1.3574 | 1.3733 | 0.77702 |
| 0.2 | 1.2274 | 1.3193 | 1.3606 | 1.3764 | 0.80239 |
| 0.25 | 1.2305 | 1.3162 | 1.3511 | 1.3701 | 0.83093 |
| 0.3 | 1.221 | 1.3128 | 1.3476 | 1.3729 | 0.85409 |
| 0.35 | 1.2305 | 1.3193 | 1.3574 | 1.3796 | 0.88167 |
| 0.4 | 1.2115 | 1.3193 | 1.3606 | 1.3764 | 0.89753 |
| 0.45 | 1.2083 | 1.3193 | 1.3574 | 1.3764 | 0.90705 |
| 0.5 | 1.2337 | 1.3162 | 1.3606 | 1.3764 | 0.90387 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.2369 | 1.332 | 1.3606 | 1.3764 | 0.93242 |
| 9.65 | 1.2527 | 1.332 | 1.3701 | 1.3859 | 0.91339 |
| 9.7 | 1.2464 | 1.3289 | 1.3637 | 1.3796 | 0.88802 |
| 9.75 | 1.2337 | 1.3257 | 1.3669 | 1.3828 | 0.86265 |
| 9.8 | 1.2496 | 1.3289 | 1.3669 | 1.3859 | 0.8341 |
| 9.85 | 1.2432 | 1.332 | 1.3669 | 1.3859 | 0.80873 |
| 9.9 | 1.2464 | 1.3225 | 1.3637 | 1.3733 | 0.78653 |
| 9.95 | 1.2368 | 1.3191 | 1.3539 | 1.3697 | 0.76552 |
| 10 | 1.2463 | 1.3159 | 1.3507 | 1.3697 | 0.75603 |

Test 4: 1.0 Hz Square Wave

A 1.0 Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 1.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1798 | 1.3067 | 1.3669 | 1.3891 | 1.1893 |
| 0.1 | 1.1766 | 1.3225 | 1.3701 | 1.3923 | 1.1861 |
| 0.15 | 1.1957 | 1.3257 | 1.3733 | 1.3955 | 1.1861 |
| 0.2 | 1.183 | 1.3162 | 1.3701 | 1.3923 | 1.1861 |
| 0.25 | 1.1893 | 1.3162 | 1.3701 | 1.3955 | 1.1861 |
| 0.3 | 1.1735 | 1.313 | 1.3796 | 1.3986 | 1.1861 |
| 0.35 | 1.1893 | 1.3193 | 1.3733 | 1.3891 | 1.1861 |
| 0.4 | 1.1766 | 1.3162 | 1.3701 | 1.3923 | 1.1861 |
| 0.45 | 1.1766 | 1.3098 | 1.3701 | 1.3955 | 1.1861 |
| 0.5 | 1.1735 | 1.313 | 1.3733 | 1.3923 | 1.1861 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 9.6 | 1.1735 | 1.3193 | 1.3733 | 1.3891 | 1.1893 |
| 9.65 | 1.1766 | 1.3289 | 1.3701 | 1.3923 | 1.1893 |
| 9.7 | 1.1703 | 1.3067 | 1.3637 | 1.3828 | 1.0054 |
| 9.75 | 1.1703 | 1.3067 | 1.3542 | 1.3764 | 1.0054 |
| 9.8 | 1.1639 | 1.3035 | 1.3574 | 1.3733 | 1.0054 |
| 9.85 | 1.1639 | 1.3003 | 1.3542 | 1.3796 | 1.0054 |
| 9.9 | 1.1766 | 1.3035 | 1.3511 | 1.3733 | 1.0022 |
| 9.95 | 1.1798 | 1.3098 | 1.3574 | 1.3796 | 1.0054 |
| 10 | 1.1608 | 1.3067 | 1.3542 | 1.3796 | 1.0054 |

Test 5: 2.0 Hz Sine Wave

A 2.0 Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 2.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.2971 | 1.3574 | 1.3828 | 1.4081 | 1.1069 |
| 0.1 | 1.2971 | 1.3637 | 1.3923 | 1.4145 | 1.1513 |
| 0.15 | 1.2876 | 1.3606 | 1.3859 | 1.4113 | 1.1639 |
| 0.2 | 1.2876 | 1.3637 | 1.3923 | 1.405 | 1.1417 |
| 0.25 | 1.2876 | 1.3606 | 1.3955 | 1.4081 | 1.091 |
| 0.3 | 1.3067 | 1.3606 | 1.3891 | 1.4081 | 1.0371 |
| 0.35 | 1.278 | 1.3476 | 1.3792 | 1.395 | 0.99327 |
| 0.4 | 1.2876 | 1.3542 | 1.3764 | 1.3891 | 0.97999 |
| 0.45 | 1.2876 | 1.3542 | 1.3828 | 1.3923 | 0.99902 |
| 0.5 | 1.2908 | 1.3542 | 1.3828 | 1.3955 | 1.0402 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.3003 | 1.3606 | 1.3828 | 1.4081 | 1.0434 |
| 4.65 | 1.294 | 1.3606 | 1.3859 | 1.4113 | 1.0973 |
| 4.7 | 1.3098 | 1.3637 | 1.3955 | 1.4081 | 1.1449 |
| 4.75 | 1.3003 | 1.3637 | 1.3955 | 1.4113 | 1.1671 |
| 4.8 | 1.3067 | 1.3733 | 1.3986 | 1.4113 | 1.1513 |
| 4.85 | 1.3003 | 1.3701 | 1.4018 | 1.4177 | 1.1037 |
| 4.9 | 1.3193 | 1.3733 | 1.3955 | 1.4081 | 1.0466 |
| 4.95 | 1.294 | 1.3574 | 1.3891 | 1.405 | 1.0054 |
| 5 | 1.2971 | 1.3606 | 1.3828 | 1.405 | 0.98316 |

Test 6: 2.0 Hz Square Wave

A 2.0 Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 2.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.2052 | 1.3225 | 1.3669 | 1.3923 | 1.1671 |
| 0.1 | 1.1957 | 1.3193 | 1.3669 | 1.3923 | 1.1671 |
| 0.15 | 1.2052 | 1.3223 | 1.3602 | 1.3887 | 1.1641 |
| 0.2 | 1.2115 | 1.3225 | 1.3669 | 1.3891 | 1.1671 |
| 0.25 | 1.1957 | 1.3225 | 1.3606 | 1.3764 | 0.98316 |
| 0.3 | 1.1798 | 1.3035 | 1.3574 | 1.3733 | 0.98316 |
| 0.35 | 1.1766 | 1.3067 | 1.3511 | 1.3669 | 0.98316 |
| 0.4 | 1.183 | 1.3035 | 1.3606 | 1.3764 | 0.98316 |
| 0.45 | 1.1893 | 1.3035 | 1.3542 | 1.3701 | 0.98316 |
| 0.5 | 1.183 | 1.313 | 1.3574 | 1.3891 | 1.1671 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1893 | 1.3193 | 1.3701 | 1.3891 | 1.1703 |
| 4.65 | 1.202 | 1.3193 | 1.3701 | 1.3891 | 1.1703 |
| 4.7 | 1.1893 | 1.3225 | 1.3701 | 1.3891 | 1.1703 |
| 4.75 | 1.1957 | 1.3193 | 1.3701 | 1.3923 | 1.1703 |
| 4.8 | 1.1925 | 1.313 | 1.3669 | 1.3796 | 0.98633 |
| 4.85 | 1.1798 | 1.3067 | 1.3574 | 1.3828 | 0.98633 |
| 4.9 | 1.1798 | 1.313 | 1.3637 | 1.3796 | 0.98633 |
| 4.95 | 1.1766 | 1.3067 | 1.3542 | 1.3701 | 0.98633 |
| 5 | 1.1767 | 1.3064 | 1.3476 | 1.3665 | 0.98379 |

Test 7: 4.0 Hz Sine Wave

A 4.0 Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 4.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.2083 | 1.313 | 1.3701 | 1.3859 | 1.1322 |
| 0.1 | 1.2147 | 1.3193 | 1.3669 | 1.3859 | 1.0402 |
| 0.15 | 1.1957 | 1.3067 | 1.3511 | 1.3669 | 0.9673 |
| 0.2 | 1.202 | 1.3162 | 1.3574 | 1.3733 | 1.0149 |
| 0.25 | 1.2115 | 1.313 | 1.3701 | 1.3796 | 1.1132 |
| 0.3 | 1.2179 | 1.3225 | 1.3415 | 1.3796 | 1.1417 |
| 0.35 | 1.2021 | 1.3096 | 1.3476 | 1.3729 | 1.0439 |
| 0.4 | 1.2179 | 1.3225 | 1.3606 | 1.3796 | 0.97048 |
| 0.45 | 1.1988 | 1.313 | 1.3542 | 1.3733 | 1.0022 |
| 0.5 | 1.2147 | 1.3193 | 1.3637 | 1.3828 | 1.1037 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.221 | 1.3257 | 1.3637 | 1.3859 | 1.1322 |
| 2.15 | 1.2147 | 1.3257 | 1.3637 | 1.3828 | 1.0339 |
| 2.2 | 1.2083 | 1.313 | 1.3542 | 1.3733 | 0.9673 |
| 2.25 | 1.2147 | 1.3098 | 1.3606 | 1.3796 | 1.018 |
| 2.3 | 1.2052 | 1.3225 | 1.3701 | 1.3891 | 1.1227 |
| 2.35 | 1.221 | 1.3225 | 1.3701 | 1.3891 | 1.1417 |
| 2.4 | 1.202 | 1.3098 | 1.3606 | 1.3859 | 1.0466 |
| 2.45 | 1.1925 | 1.3067 | 1.3542 | 1.3701 | 0.97048 |
| 2.5 | 1.1988 | 1.3098 | 1.3606 | 1.3764 | 1.0085 |

Test 8: 4.0 Hz Square Wave

A 4.0 Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 4.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.202 | 1.313 | 1.3733 | 1.3923 | 1.202 |
| 0.1 | 1.1957 | 1.313 | 1.3637 | 1.3859 | 1.018 |
| 0.15 | 1.1766 | 1.3162 | 1.3574 | 1.3764 | 1.018 |
| 0.2 | 1.1893 | 1.3067 | 1.3606 | 1.3828 | 1.018 |
| 0.25 | 1.183 | 1.3225 | 1.3701 | 1.3955 | 1.202 |
| 0.3 | 1.1988 | 1.3225 | 1.3669 | 1.3955 | 1.202 |
| 0.35 | 1.1861 | 1.3193 | 1.3637 | 1.3923 | 1.202 |
| 0.4 | 1.1988 | 1.3162 | 1.3669 | 1.3796 | 1.0149 |
| 0.45 | 1.183 | 1.313 | 1.3606 | 1.3828 | 1.018 |
| 0.5 | 1.1988 | 1.3225 | 1.3733 | 1.3891 | 1.202 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1798 | 1.313 | 1.3606 | 1.3764 | 1.018 |
| 2.15 | 1.183 | 1.3098 | 1.3542 | 1.3764 | 1.018 |
| 2.2 | 1.1798 | 1.3098 | 1.3669 | 1.3828 | 1.018 |
| 2.25 | 1.1925 | 1.3162 | 1.3637 | 1.3923 | 1.202 |
| 2.3 | 1.1925 | 1.3225 | 1.3637 | 1.3859 | 1.202 |
| 2.35 | 1.2052 | 1.3162 | 1.3701 | 1.3828 | 1.018 |
| 2.4 | 1.1798 | 1.313 | 1.3637 | 1.3764 | 1.0149 |
| 2.45 | 1.1798 | 1.3098 | 1.3574 | 1.3764 | 1.018 |
| 2.5 | 1.202 | 1.3193 | 1.3764 | 1.3955 | 1.202 |

Test 9: 5.0 Hz Sine Wave

A 5.0 Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 5.0 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.2052 | 1.3098 | 1.3606 | 1.3859 | 0.98951 |
| 0.1 | 1.1988 | 1.3098 | 1.3511 | 1.3764 | 0.98316 |
| 0.15 | 1.202 | 1.313 | 1.3606 | 1.3923 | 1.11 |
| 0.2 | 1.202 | 1.3225 | 1.3637 | 1.3859 | 1.1259 |
| 0.25 | 1.202 | 1.3193 | 1.3637 | 1.3764 | 0.99902 |
| 0.3 | 1.1988 | 1.3035 | 1.3574 | 1.3764 | 0.97682 |
| 0.35 | 1.1957 | 1.313 | 1.3606 | 1.3828 | 1.0973 |
| 0.4 | 1.1957 | 1.313 | 1.3606 | 1.3859 | 1.1322 |
| 0.45 | 1.2083 | 1.3162 | 1.3669 | 1.3828 | 1.0149 |
| 0.5 | 1.1957 | 1.3098 | 1.3574 | 1.3733 | 0.97048 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.202 | 1.3098 | 1.3606 | 1.3796 | 0.98951 |
| 2.15 | 1.183 | 1.3067 | 1.3606 | 1.3733 | 0.98951 |
| 2.2 | 1.202 | 1.3162 | 1.3669 | 1.3828 | 1.1164 |
| 2.25 | 1.2083 | 1.3193 | 1.3669 | 1.3859 | 1.1227 |
| 2.3 | 1.1861 | 1.3162 | 1.3606 | 1.3733 | 0.99902 |
| 2.35 | 1.202 | 1.313 | 1.3637 | 1.3764 | 0.98316 |
| 2.4 | 1.1988 | 1.3098 | 1.3574 | 1.3891 | 1.1037 |
| 2.45 | 1.1925 | 1.3193 | 1.3606 | 1.3859 | 1.1322 |
| 2.5 | 1.1893 | 1.313 | 1.3574 | 1.3733 | 1.0149 |

Test 10: 5.0 Hz Square Wave

A 5.0 Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 5.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1861 | 1.313 | 1.3542 | 0.92608 | 1.0022 |
| 0.1 | 1.1608 | 1.2908 | 1.3384 | 0.92925 | 1.0022 |
| 0.15 | 1.1798 | 1.3067 | 1.3542 | 0.9451 | 1.1893 |
| 0.2 | 1.1798 | 1.3035 | 1.3542 | 0.93559 | 1.1861 |
| 0.25 | 1.1925 | 1.3098 | 1.3637 | 0.93876 | 1.1861 |
| 0.3 | 1.1798 | 1.3003 | 1.3542 | 0.92925 | 1.0022 |
| 0.35 | 1.1608 | 1.2876 | 1.3447 | 0.92925 | 1.0022 |
| 0.4 | 1.1671 | 1.3098 | 1.3574 | 0.93559 | 1.1861 |
| 0.45 | 1.1861 | 1.3098 | 1.3542 | 0.93559 | 1.1861 |
| 0.5 | 1.1671 | 1.313 | 1.3574 | 0.92608 | 1.0022 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1766 | 1.3067 | 1.3511 | 0.95779 | 1.0022 |
| 2.15 | 1.1639 | 1.2876 | 1.3415 | 0.96413 | 1.0022 |
| 2.2 | 1.1671 | 1.3003 | 1.3447 | 0.97365 | 1.1861 |
| 2.25 | 1.1735 | 1.3098 | 1.3637 | 0.97365 | 1.1861 |
| 2.3 | 1.1893 | 1.3098 | 1.3542 | 0.9673 | 0.99902 |
| 2.35 | 1.1671 | 1.313 | 1.3542 | 0.96096 | 1.0022 |
| 2.4 | 1.1639 | 1.2971 | 1.3479 | 0.96096 | 1.0022 |
| 2.45 | 1.183 | 1.3035 | 1.3574 | 0.97365 | 1.1861 |
| 2.5 | 1.221 | 1.3162 | 1.3574 | 0.96096 | 1.1861 |

Test 11: 7.5 Hz Sine Wave

A 7.5 Hz Sine Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 7.5 Hz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1893 | 1.3098 | 1.3669 | 1.3986 | 1.1132 |
| 0.1 | 1.1798 | 1.3162 | 1.3669 | 1.3859 | 0.99585 |
| 0.15 | 1.1735 | 1.313 | 1.3637 | 1.3859 | 0.97999 |
| 0.2 | 1.1861 | 1.3193 | 1.3764 | 1.4018 | 1.1195 |
| 0.25 | 1.1639 | 1.313 | 1.3669 | 1.3859 | 0.96096 |
| 0.3 | 1.1608 | 1.3098 | 1.3637 | 1.3891 | 1.0307 |
| 0.35 | 1.1671 | 1.3162 | 1.3669 | 1.4018 | 1.0878 |
| 0.4 | 1.1608 | 1.3098 | 1.3637 | 1.3891 | 0.93559 |
| 0.45 | 1.1671 | 1.313 | 1.3669 | 1.3923 | 1.0846 |
| 0.5 | 1.1766 | 1.3193 | 1.3701 | 1.3923 | 1.0371 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1576 | 1.2971 | 1.3701 | 1.3955 | 1.1195 |
| 2.15 | 1.1481 | 1.2971 | 1.3637 | 1.3891 | 0.99585 |
| 2.2 | 1.1608 | 1.2971 | 1.3637 | 1.3828 | 0.98951 |
| 2.25 | 1.1639 | 1.3098 | 1.3701 | 1.3923 | 1.1227 |
| 2.3 | 1.1513 | 1.3035 | 1.3637 | 1.3828 | 0.95462 |
| 2.35 | 1.1513 | 1.3003 | 1.3669 | 1.3891 | 1.0434 |
| 2.4 | 1.2052 | 1.3193 | 1.3669 | 1.3859 | 1.0846 |
| 2.45 | 1.1608 | 1.3067 | 1.3511 | 1.3733 | 0.94193 |
| 2.5 | 1.1576 | 1.3003 | 1.3606 | 1.3859 | 1.0973 |

Test 12: 7.5 Hz Square Wave

A 7.5 Hz Square Wave was applied from the function generator with 500mV peak to peak and no DC offset. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 7.5 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (s) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1988 | 1.3257 | 1.3669 | 1.3955 | 1.1893 |
| 0.1 | 1.1861 | 1.3225 | 1.3701 | 1.3955 | 1.1861 |
| 0.15 | 1.1861 | 1.3098 | 1.3637 | 1.3828 | 1.0022 |
| 0.2 | 1.1925 | 1.3162 | 1.3637 | 1.3955 | 1.1861 |
| 0.25 | 1.2052 | 1.313 | 1.3669 | 1.3796 | 1.0022 |
| 0.3 | 1.183 | 1.3098 | 1.3542 | 1.3796 | 1.0022 |
| 0.35 | 1.2052 | 1.3193 | 1.3701 | 1.3859 | 1.1893 |
| 0.4 | 1.1861 | 1.313 | 1.3606 | 1.3796 | 1.0022 |
| 0.45 | 1.202 | 1.3225 | 1.3669 | 1.3891 | 1.1893 |
| 0.5 | 1.1988 | 1.313 | 1.3701 | 1.3923 | 1.1893 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 2.1 | 1.1957 | 1.3225 | 1.3637 | 1.3955 | 1.1893 |
| 2.15 | 1.1925 | 1.3225 | 1.3764 | 1.3923 | 1.1893 |
| 2.2 | 1.1925 | 1.3162 | 1.3637 | 1.3828 | 1.0054 |
| 2.25 | 1.202 | 1.3225 | 1.3637 | 1.3986 | 1.1893 |
| 2.3 | 1.1925 | 1.3193 | 1.3637 | 1.3796 | 1.0054 |
| 2.35 | 1.1988 | 1.313 | 1.3606 | 1.3796 | 1.0054 |
| 2.4 | 1.1957 | 1.3162 | 1.3701 | 1.3859 | 1.1893 |
| 2.45 | 1.1925 | 1.3098 | 1.3606 | 1.3733 | 1.0054 |
| 2.5 | 1.1925 | 1.3193 | 1.3733 | 1.3923 | 1.1893 |

Conclusion:

The Full System Reconstruction Test was deemed successful.  The EKG sensor was able to change the signal and then the MICA2 mote was able to transmit the signal from the EKG sensor output.  The Moteview Software was able to collect the data and store it in a csv file.  The data was later graphed and tabulated for visualization purposes.  The Full System was able to view the data and reconstruct a signal successfully up to 4.0Hz.  This is successful due to the EKG signal only beating at a rate of roughly 1.8Hz.

1-Hop Full System Test

The 1-Hop Full System Test was a continuation of the Full System Reconstruction Test.  This test consisted of the EKG sensor being attached to a MICA2 Mote. The data from the MICA2 mote  then needed to be wirelessly transmitted  to the base station.  The setup for this test can be seen below in the figure.  A function generator was used to apply both Sine and square waves.  Frequencies of 100mHz, 300mHz, and 500mHz were used.  The RF mote supplied the necessary 3V of power needed to operate the EKG sensor.  The output of the EKG sensor was sent to ADC6 of the MDA100CB data acquisition board connected to the MICA2 mote.

Test 1: 100mHz Sine Wave

A  100mHz Sine Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 100mHz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.6908 | 2.5251 | 2.3563 | 2.1907 | 1.2856 |
| 0.6 | 2.6908 | 2.5312 | 2.3656 | 2.2029 | 1.2641 |
| 0.9 | 2.6846 | 2.5343 | 2.3686 | 2.2091 | 1.2457 |
| 1.2 | 2.6785 | 2.5312 | 2.3656 | 2.2152 | 1.2273 |
| 1.5 | 2.6908 | 2.5374 | 2.3686 | 2.2183 | 1.215 |
| 1.8 | 2.6877 | 2.5343 | 2.3656 | 2.2152 | 1.2058 |
| 2.1 | 2.6754 | 2.5343 | 2.3686 | 2.2213 | 1.1997 |
| 2.4 | 2.6969 | 2.5374 | 2.3748 | 2.2213 | 1.1966 |
| 2.7 | 2.6938 | 2.5374 | 2.3717 | 2.2213 | 1.1935 |
| 3 | 2.6871 | 2.5341 | 2.3658 | 2.2189 | 1.1936 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 37.8 | 2.6657 | 2.4759 | 2.2739 | 2.0964 | 1.4568 |
| 38.1 | 2.6724 | 2.5067 | 2.3165 | 2.1477 | 1.4758 |
| 38.4 | 2.6908 | 2.522 | 2.3441 | 2.1784 | 1.4819 |
| 38.7 | 2.6877 | 2.5282 | 2.3594 | 2.1968 | 1.4789 |
| 39 | 2.6871 | 2.5249 | 2.3535 | 2.2036 | 1.4568 |
| 39.3 | 2.6938 | 2.5374 | 2.3717 | 2.2152 | 1.4359 |
| 39.6 | 2.6718 | 2.531 | 2.3596 | 2.2127 | 1.4048 |
| 39.9 | 2.6687 | 2.531 | 2.3627 | 2.2189 | 1.3711 |
| 40.2 | 2.6908 | 2.5404 | 2.3686 | 2.2213 | 1.3439 |

Test 2: 100mHz Square Wave

A  100mHz Square Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 100mHz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.659 | 2.5186 | 2.3568 | 2.2133 | 1.0746 |
| 0.6 | 2.6651 | 2.5278 | 2.3629 | 2.2164 | 1.0746 |
| 0.9 | 2.6749 | 2.528 | 2.3658 | 2.2189 | 1.0773 |
| 1.2 | 1.6027 | 1.8806 | 1.9844 | 1.9874 | 1.0716 |
| 1.5 | 2.4087 | 2.1584 | 2.0301 | 1.9569 | 1.0746 |
| 1.8 | 2.5522 | 2.3171 | 2.1278 | 1.9966 | 1.0746 |
| 2.1 | 1.5722 | 1.7859 | 1.8348 | 1.8256 | 1.0716 |
| 2.4 | 2.3904 | 2.1156 | 1.9508 | 1.8531 | 1.0716 |
| 2.7 | 2.5552 | 2.3049 | 2.0912 | 1.9325 | 1.0716 |
| 3 | 2.632 | 2.4178 | 2.2036 | 2.023 | 1.0742 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 37.8 | 2.6682 | 2.5247 | 2.366 | 2.2225 | 1.4013 |
| 38.1 | 2.6682 | 2.5217 | 2.3629 | 2.2194 | 1.3982 |
| 38.4 | 2.6774 | 2.5217 | 2.366 | 2.2194 | 1.3982 |
| 38.7 | 2.6774 | 2.5217 | 2.366 | 2.2194 | 1.3982 |
| 39 | 2.6835 | 2.5278 | 2.366 | 2.2194 | 1.3982 |
| 39.3 | 2.6774 | 2.5308 | 2.366 | 2.2225 | 1.3982 |
| 39.6 | 1.6058 | 1.8897 | 1.9874 | 1.9935 | 1.3982 |
| 39.9 | 2.3964 | 2.1576 | 2.0352 | 1.9618 | 1.4017 |
| 40.2 | 2.5613 | 2.3171 | 2.1217 | 1.9752 | 1.0654 |

Test 3: 300mHz Sine Wave

A  300mHz Sine Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 300mHz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.6646 | 2.5093 | 2.3388 | 2.1865 | 1.3491 |
| 0.6 | 2.6519 | 2.497 | 2.336 | 2.1902 | 1.3761 |
| 0.9 | 2.6494 | 2.5063 | 2.3418 | 2.1926 | 1.3125 |
| 1.2 | 2.6428 | 2.494 | 2.336 | 2.1841 | 1.209 |
| 1.5 | 2.655 | 2.5 | 2.339 | 2.1841 | 1.124 |
| 1.8 | 2.6398 | 2.5031 | 2.339 | 2.1841 | 1.0662 |
| 2.1 | 2.655 | 2.5031 | 2.3421 | 2.1871 | 1.0389 |
| 2.4 | 2.6494 | 2.5032 | 2.3418 | 2.1956 | 1.0415 |
| 2.7 | 2.655 | 2.5 | 2.339 | 2.1963 | 1.0662 |
| 3 | 2.6524 | 2.5093 | 2.3479 | 2.2048 | 1.1298 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 7.8 | 1.5918 | 1.7983 | 1.8469 | 1.8317 | 1.2849 |
| 8.1 | 2.3876 | 2.1082 | 1.9472 | 1.8439 | 1.1847 |
| 8.4 | 2.5367 | 2.2961 | 2.086 | 1.9246 | 1.1085 |
| 8.7 | 2.6124 | 2.3968 | 2.1841 | 1.9988 | 1.0571 |
| 9 | 2.6337 | 2.4514 | 2.2509 | 2.0656 | 1.0359 |
| 9.3 | 2.6337 | 2.4757 | 2.2935 | 2.1173 | 1.0419 |
| 9.6 | 2.6458 | 2.494 | 2.3178 | 2.1477 | 1.0784 |
| 9.9 | 2.655 | 2.494 | 2.3299 | 2.172 | 1.1422 |
| 10.2 | 2.6489 | 2.5 | 2.336 | 2.1871 | 1.2363 |

Test 4: 300mHz Square Wave

A  300mHz Square Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 300mHz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.3753 | 2.0982 | 1.9368 | 1.8424 | 1.3856 |
| 0.6 | 2.5337 | 2.287 | 2.0769 | 1.9216 | 1.3856 |
| 0.9 | 2.6159 | 2.3936 | 2.1774 | 1.9886 | 1.0506 |
| 1.2 | 2.6342 | 2.4514 | 2.2474 | 2.0616 | 1.0506 |
| 1.5 | 2.659 | 2.4942 | 2.3049 | 2.1248 | 1.0532 |
| 1.8 | 2.6676 | 2.5032 | 2.3205 | 2.153 | 1.0506 |
| 2.1 | 2.6585 | 2.5093 | 2.3388 | 2.1743 | 1.0506 |
| 2.4 | 2.6737 | 2.5123 | 2.3509 | 2.2017 | 1.3308 |
| 2.7 | 2.6585 | 2.5184 | 2.357 | 2.2139 | 1.3856 |
| 3 | 2.6743 | 2.5217 | 2.3568 | 2.2164 | 1.389 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 7.8 | 2.6804 | 2.5247 | 2.3568 | 2.1981 | 1.0532 |
| 8.1 | 2.6585 | 2.5154 | 2.354 | 2.2048 | 1.0506 |
| 8.4 | 2.6774 | 2.5217 | 2.3599 | 2.2072 | 1.0532 |
| 8.7 | 1.618 | 1.8928 | 1.9874 | 1.9905 | 1.0532 |
| 9 | 2.4088 | 2.153 | 2.0312 | 1.952 | 1.0506 |
| 9.3 | 2.5489 | 2.3174 | 2.1317 | 2.0129 | 1.3856 |
| 9.6 | 1.5592 | 1.7784 | 1.8363 | 1.8393 | 1.3856 |
| 9.9 | 2.3782 | 2.1095 | 1.9508 | 1.8561 | 1.389 |
| 10.2 | 2.5461 | 2.2957 | 2.0851 | 1.9355 | 1.389 |

Test 5: 500mHz Sine Wave

A  500mHz Sine Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 500mHz Sine Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.3519 | 2.0896 | 1.9328 | 1.8272 | 1.0915 |
| 0.6 | 2.5117 | 2.2675 | 2.0624 | 1.8996 | 1.0192 |
| 0.9 | 2.5931 | 2.379 | 2.168 | 1.9901 | 1.0282 |
| 1.2 | 2.6263 | 2.4363 | 2.2403 | 2.0624 | 1.1247 |
| 1.5 | 2.6384 | 2.4695 | 2.2856 | 2.1197 | 1.2905 |
| 1.8 | 2.6414 | 2.4816 | 2.3097 | 2.1529 | 1.3689 |
| 2.1 | 2.6293 | 2.4846 | 2.3157 | 2.1559 | 1.2121 |
| 2.4 | 2.6293 | 2.4876 | 2.3187 | 2.1589 | 1.0734 |
| 2.7 | 2.6293 | 2.4846 | 2.3218 | 2.168 | 1.0161 |
| 3 | 2.6444 | 2.4936 | 2.3338 | 2.18 | 1.0403 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 7.8 | 2.6353 | 2.4936 | 2.3338 | 2.2072 | 1.3659 |
| 8.1 | 2.6323 | 2.4966 | 2.3368 | 2.1951 | 1.2996 |
| 8.4 | 2.6444 | 2.4936 | 2.3338 | 2.1861 | 1.1277 |
| 8.7 | 1.5981 | 1.8695 | 1.9599 | 1.9629 | 1.0312 |
| 9 | 2.3579 | 2.1258 | 2.0051 | 1.9268 | 1.0161 |
| 9.3 | 2.5268 | 2.2916 | 2.1047 | 1.972 | 1.0855 |
| 9.6 | 1.5529 | 1.7669 | 1.8242 | 1.8182 | 1.2332 |
| 9.9 | 2.3429 | 2.0805 | 1.9268 | 1.8423 | 1.375 |
| 10.2 | 2.5117 | 2.2705 | 2.0624 | 1.9117 | 1.2664 |

Test 6: 500mHz Square Wave

A  500mHz Square Wave was applied from the function generator with 400mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 500mHz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.3 | 2.6519 | 2.5152 | 2.3481 | 2.2084 | 1.3761 |
| 0.6 | 2.6489 | 2.5061 | 2.3421 | 2.1811 | 1.0298 |
| 0.9 | 2.6519 | 2.5061 | 2.3421 | 2.1902 | 1.0298 |
| 1.2 | 2.661 | 2.5122 | 2.3481 | 2.1993 | 1.0298 |
| 1.5 | 2.6484 | 2.5029 | 2.3484 | 2.212 | 1.3727 |
| 1.8 | 2.6519 | 2.5152 | 2.3542 | 2.2145 | 1.3761 |
| 2.1 | 2.6519 | 2.5091 | 2.3512 | 2.2084 | 1.3761 |
| 2.4 | 2.6671 | 2.5122 | 2.3512 | 2.2114 | 1.3761 |
| 2.7 | 2.6423 | 2.4969 | 2.3363 | 2.1787 | 1.0272 |
| 3 | 2.6423 | 2.4999 | 2.3393 | 2.1848 | 1.0272 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 7.8 | 2.6544 | 2.4938 | 2.3211 | 2.1666 | 1.3727 |
| 8.1 | 2.6423 | 2.4969 | 2.3302 | 2.1787 | 1.3727 |
| 8.4 | 2.6458 | 2.5 | 2.336 | 2.1871 | 1.3761 |
| 8.7 | 2.6641 | 2.5061 | 2.336 | 2.175 | 1.0298 |
| 9 | 2.6519 | 2.5 | 2.339 | 2.1841 | 1.0298 |
| 9.3 | 2.6635 | 2.506 | 2.3393 | 2.1878 | 1.0272 |
| 9.6 | 2.6454 | 2.506 | 2.3393 | 2.1939 | 1.0272 |
| 9.9 | 2.6671 | 2.5152 | 2.3542 | 2.2206 | 1.3761 |
| 10.2 | 2.6575 | 2.509 | 2.3454 | 2.209 | 1.3727 |

Conclusion:

From the 1-Hop Full System test it was verified that the full system worked with 1-Hop.  However due to coding complications only a frequency of 500mHz could be applied with the ability to reconstruct the signal at the base station.  However the data is as expected from transmitting MICA2.  This test was deemed successful despite not being able to reach a higher frequency.

Heart Beat Test

EKG Sensor

Base Station

RF Mote

USB

Sync

Output

Waveform Generator

The Heart Beat Test were performed by applying a signal from a function generation to the EKG sensor board.  The output signal was then linked to ADC6 of the MDA100CB data acquisition board.  The data acquisition board was attached to the MICA2 mote which was directly connected to the MIB520 base station.  The figure below shows the setup of the test. The MICA2 Mote supplies the power for the whole system.  Sine waves and square waves were used at different frequencies to test the ability of the MICA2 mote and Moteview software to reconstruct the signal through the EKG sensor board.  Frequencies of 1.67Hz, 2Hz, and 3Hz were used to determine the systems operation. These frequencies were chosen due to the relationship they have to a normal heartbeat.

Test 1: 1.67Hz Square Wave

A 1.67 Hz Square Wave was applied from the function generator with 50mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 1.67 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1893 | 1.313 | 1.3574 | 1.3828 | 1.2749 |
| 0.1 | 1.1925 | 1.3098 | 1.3542 | 1.3828 | 1.2781 |
| 0.15 | 1.1893 | 1.3162 | 1.3606 | 1.3828 | 1.2781 |
| 0.2 | 1.1798 | 1.3098 | 1.3574 | 1.3828 | 1.2781 |
| 0.25 | 1.1893 | 1.313 | 1.3574 | 1.3796 | 1.2527 |
| 0.3 | 1.1988 | 1.3067 | 1.3542 | 1.3796 | 1.2527 |
| 0.35 | 1.1861 | 1.3035 | 1.3606 | 1.3796 | 1.2527 |
| 0.4 | 1.2052 | 1.3098 | 1.3637 | 1.3796 | 1.2527 |
| 0.45 | 1.1893 | 1.3035 | 1.3574 | 1.3733 | 1.2527 |
| 0.5 | 1.1988 | 1.3098 | 1.3637 | 1.3764 | 1.2527 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1957 | 1.3162 | 1.3701 | 1.3859 | 1.2527 |
| 4.65 | 1.1766 | 1.3257 | 1.3701 | 1.3764 | 1.2527 |
| 4.7 | 1.1861 | 1.3193 | 1.3606 | 1.3796 | 1.2527 |
| 4.75 | 1.1798 | 1.3162 | 1.3637 | 1.3828 | 1.2527 |
| 4.8 | 1.202 | 1.313 | 1.3669 | 1.3828 | 1.2527 |
| 4.85 | 1.1925 | 1.3067 | 1.3669 | 1.3891 | 1.2813 |
| 4.9 | 1.1957 | 1.3162 | 1.3637 | 1.3859 | 1.2781 |
| 4.95 | 1.1766 | 1.313 | 1.3637 | 1.3891 | 1.2781 |
| 5 | 1.1862 | 1.3128 | 1.3571 | 1.376 | 1.2748 |

Test 2: 2.0 Hz Square Wave

A 2.0 Hz Square Wave was applied from the function generator with 50mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 2.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1735 | 1.313 | 1.3637 | 1.3859 | 1.2559 |
| 0.1 | 1.1639 | 1.313 | 1.3606 | 1.3859 | 1.2559 |
| 0.15 | 1.183 | 1.3162 | 1.3733 | 1.3891 | 1.2559 |
| 0.2 | 1.1766 | 1.3162 | 1.3764 | 1.3955 | 1.2813 |
| 0.25 | 1.183 | 1.3225 | 1.3733 | 1.3923 | 1.2813 |
| 0.3 | 1.1703 | 1.3193 | 1.3764 | 1.3955 | 1.2813 |
| 0.35 | 1.183 | 1.3193 | 1.3701 | 1.3923 | 1.2781 |
| 0.4 | 1.1609 | 1.3096 | 1.3634 | 1.3855 | 1.278 |
| 0.45 | 1.1671 | 1.313 | 1.3764 | 1.3891 | 1.2527 |
| 0.5 | 1.1608 | 1.3193 | 1.3701 | 1.3891 | 1.2559 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1671 | 1.313 | 1.3701 | 1.3891 | 1.2527 |
| 4.65 | 1.1798 | 1.313 | 1.3669 | 1.3891 | 1.2527 |
| 4.7 | 1.1735 | 1.3162 | 1.3701 | 1.3923 | 1.2845 |
| 4.75 | 1.1639 | 1.313 | 1.3733 | 1.3955 | 1.2845 |
| 4.8 | 1.1671 | 1.313 | 1.3764 | 1.3955 | 1.2845 |
| 4.85 | 1.1703 | 1.3193 | 1.3733 | 1.3923 | 1.2876 |
| 4.9 | 1.1671 | 1.3257 | 1.3733 | 1.3891 | 1.2876 |
| 4.95 | 1.1639 | 1.3225 | 1.3733 | 1.3923 | 1.2559 |
| 5 | 1.183 | 1.313 | 1.3733 | 1.3923 | 1.2559 |

Test 3: 3.0 Hz Square Wave

A 3.0 Hz Square Wave was applied from the function generator with 50mV peak to peak and a DC offset of 400mV. The data was exported from Moteview and graphed in Excel.  The figure below shows the reconstructed graph of the 3.0 Hz Square Wave. The table below shows the raw data from ADC6.

| Time (sec) | adc2 [V] | adc3 [V] | adc4 [V] | adc5 [V] | adc6 [V] |
| --- | --- | --- | --- | --- | --- |
| 0.05 | 1.1639 | 1.3225 | 1.3764 | 1.3986 | 1.2559 |
| 0.1 | 1.1671 | 1.3193 | 1.3733 | 1.3955 | 1.2591 |
| 0.15 | 1.1671 | 1.313 | 1.3701 | 1.3986 | 1.2591 |
| 0.2 | 1.1544 | 1.3098 | 1.3669 | 1.3923 | 1.2845 |
| 0.25 | 1.1608 | 1.3098 | 1.3733 | 1.3955 | 1.2845 |
| 0.3 | 1.1735 | 1.3067 | 1.3606 | 1.3891 | 1.2845 |
| 0.35 | 1.1703 | 1.3162 | 1.3733 | 1.3955 | 1.2813 |
| 0.4 | 1.1481 | 1.3098 | 1.3701 | 1.3923 | 1.2813 |
| 0.45 | 1.1544 | 1.3162 | 1.3701 | 1.3986 | 1.2527 |
| 0.5 | 1.1514 | 1.3159 | 1.3634 | 1.3855 | 1.2495 |
| . | . | . | . | . | . |
| . | . | . | . | . | . |
| 4.6 | 1.1544 | 1.3098 | 1.3733 | 1.3923 | 1.2527 |
| 4.65 | 1.1576 | 1.313 | 1.3733 | 1.3923 | 1.2527 |
| 4.7 | 1.1546 | 1.3096 | 1.3665 | 1.3855 | 1.2463 |
| 4.75 | 1.1608 | 1.3162 | 1.3733 | 1.3923 | 1.2527 |
| 4.8 | 1.1576 | 1.3098 | 1.3701 | 1.3923 | 1.2813 |
| 4.85 | 1.1735 | 1.3098 | 1.3733 | 1.3891 | 1.2845 |
| 4.9 | 1.1671 | 1.3098 | 1.3733 | 1.3955 | 1.2845 |
| 4.95 | 1.1513 | 1.3067 | 1.3637 | 1.3859 | 1.2876 |
| 5 | 1.1671 | 1.3098 | 1.3701 | 1.3891 | 1.2876 |

Conclusion:

The Heart Rate Test was deemed successful.  The EKG sensor was able to change the signal and then the MICA2 mote was able to transmit the signal from the EKG sensor output.  The Moteview Software was able to collect the data and store it in a csv file.  The data was later graphed and tabulated for visualization purposes.  This test is successful due to the EKG being viewed at these different frequencies at which heart would beat.

Budget

The budget for our project is broken down into two main categories: Capstone Budget and Out of Pocket expenditures.  The table and pie chart below shows the breakdown of the total money allotted to our project.

| Category | Amount ($) |
| --- | --- |
| Out of Pocket Expenditures | $49.93 |
| Capstone Budget Spent | $427.82 |
| Capstone Budget Unspent | $72.18 |
| Total | $549.93 |

The department allotted $500 toward our EKG Project.  We spent $427.82 out of the $500 leaving $72.18 unspent.  The table and pie chart below shows a breakdown of how the Capstone budget money was spent.

| Category | Amount ($) |
| --- | --- |
| Cables & Connectors | $111.75 |
| Advanced Circuit (PCB) | $237.99 |
| Components | $78.08 |
| Unspent | $72.18 |
| Total | $500.00 |

Our Out of Pocket Expenditures was made up of two categories: shipping components and board to be soldered and electrodes.  The table and pie chart below shows our Out of Pocket expenditures.

| Category | Amount ($) |
| --- | --- |
| Shipping | $40.53 |
| Electrodes | $9.40 |
| Total | $49.93 |

Our Order Request Forms and Invoices for our Capstone Budget can be seen below:

Order Request Form

Department of Electrical and Computer Engineering

ECE Staff Only

Date: __________8/24/10____________		Pcard_______   Requisition ________

| Company Name: | Digikey |
| --- | --- |
| Address: | Digi-Key Corporation 701 Brooks Avenue South Thief River Falls, MN 56701 USA |
| Website URL: | www.digikey.com |
| Phone Number: | 1-800-344-4539 |
| Fax Number: | 218-681-3380 |
| Contact Person: |  |
| Vendor Information Form Needed? |  |

| Account Number to be Used: | Dr. Hu |
| --- | --- |
| End date for grant/contract (if purchase is to be made from grant/contract account) |  |
| Requesting Person: Phone #: Fax #: Email Address: |  |
|  | n/a |

Please print clearly.

| Quantity | Part # | Description | Price | Total |
| --- | --- | --- | --- | --- |
| 5 | INA321ECT-ND | INA321 OP AMP | $5.10 | $25.50 |
| 20 | 541-1.0MACT-ND | 1MOhms Resistor, 5% | $0.077 | $1.54 |
| Shipping and Handling (to be included) Total |  |  |  |  |
|  |  |  |  | $27.04 |

Signature Approval: ________________________

(only authorized investigators may sign for grant/contract purchases)

Order Request Form

Department of Electrical and Computer Engineering

ECE Staff Only

Date: ______8/23/10_______________   	Pcard_______   Requisition ________

| Company Name: | Mouser Electronics |
| --- | --- |
| Address: | Mouser Electronics 1000 North Main Street Mansfield, TX 76063 |
| Website URL: | www.mouser.com |
| Phone Number: | (800) 298-5076 |
| Fax Number: | (817) 804-3899 |
| Contact Person: | Larry Rosen |
| Vendor Information Form Needed? |  |

| Account Number to be Used: | Dr.Hu |
| --- | --- |
| End date for grant/contract (if purchase is to be made from grant/contract account) |  |
| Requesting Person: Phone #: Fax #: Email Address: |  |
|  | n/a |

Please print clearly.

| Quantity | Part # | Description | Price | Total |
| --- | --- | --- | --- | --- |
| 5 | 595-OPA4336EA/250 | OPA4336 Op-Amp made by Texas Instr. | $4.49 | $22.45 |
| 13 | 71-CRCW0805100kjneb | 100kOhms Resistor, 5% | $0.07 | $0.91 |
| 25 | 71-CRCW080547K0JNEA | 47kOhms Resistor, 5% | $0.04 | $1.00 |
| 25 | 71-CRCW0805J-2.2M-E3 | 2.2Mohms Resistor, 5% | $0.04 | $1.00 |
| 25 | 71-CRCW08053K30JNEA | 3.3kOhms Resistor, 5% | $0.04 | $1.00 |
| 25 | 71-CRCW0805J-4.7K-E3 | 4.7kOhms Resistor, 5% | $0.04 | $1.00 |
| 10 | 71-CRCW0805-806-E3 | 806 Ohms Resistor, 5% | $0.05 | $0.50 |
| 15 | 77-VJ0805Y104KXARBC | 0.1uF Capacitor, 50V | $0.15 | $2.25 |
| 10 | 80-C1206C472K1RACTM | 4.7nF Capacitor, 100V | $0.38 | $3.80 |
| 10 | 667-ECP-U1C105MA5 | 1uF Capacitor, 16V | $0.58 | $5.80 |
| 15 | 598-FCP1210C104G-G3 | 0.1uF Capacitor | $0.36 | $5.40 |
| Shipping and Handling (to be included) Total |  |  |  |  |
|  |  |  |  | $45.11 |

If this request is within 60 days of the end of a contact and a Pcard purchase is requested, I verify that sufficient funds exist in the grant/contract account for this purchase.

Signature Approval: ________________________

(only authorized investigators may sign for grant/contract purchases)

Our invoices and receipts for our Out of Pocket Expenditures can be seen below:

Your order number 0225784 has been invoiced with invoice number  0261724-IN.  Please review the detail below and note any backordered items and changes to sales tax and/or discount amounts.

Payment Method: VISA

US                                          US

_______________________________________________________________________________

Ship Via:    FEDEX GROUND                   Taxable:                     2.70

Terms:       50                             Non-Taxable:                 0.00

CREDIT CARD                    Freight:                     6.70

Customer PO: INTERNET ORDER                 Sales Tax:                   0.00

Ship Date:   11/17/2010

Order Total:                 9.40

_______________________________________________________________________________

Item Number              Qty Shipped           Unit Price       Extension

11770                             6                 0.45            2.70

ELECTRODES, SNAP ECG 3/POUCH

REPLACES 36003

References

[1]  	Hu, Fei, Yang Xiao, and Qi Hao.  “Congestion-aware, Loss-Resilient Bio-monitoring Sensor Networking for Mobile Health Applications.” IEEE Journal on Selected Areas in Communications 27.4 (2009): 450-465.

[2]  	Fulford-Jones, Thaddeus, David Malan, Steve Moulton, and Matt Welsh. “CodeBlue: An Ad Hoc Sensor Network Infrastructure for Emergency Medical Care.” Workshop on Applications of Mobile Embedded Systems, June 2004.

[3]  	Gaynor, Mark, Ed LaCombe, Steve Moulton, Austin Rowan, Matt Welsh and John Wynee. “Wireless Sensor Network Applications.” Americas Conference on Information Systems, August 2004.

[4]  	Fulford-Jones, Thaddeus, Gu-Yeon Wei, and Matt Welsh. “A Portable, Low-Power, Wireless Two-Lead EKG System.”  Proceedings of the 26th Annual International Conference of the IEEE EMBS, 1 September 2004 – 5 September 2004.  San Francisco, California.

[5]  	Clavel, Antony, Thaddeus Fulford-Jones, Konrad Lorincz, Geoffrey Mainland, David Malan, Alan Nawoj, Victor Shnayder, and Matt Welsh.  “Sensor Networks for Emergency Response: Challenges and Opportunities.” IEEE Pervasive Computing 3.4 (2004): 16-23.

[6]  	Chen, Bor-rong, Thaddeus Fulford-Jones, Victor Shnayder, and Matt Welsh. Harvard University.  “Sensor Networks for Medical Care.” 2005. 11 January 2010. <http://www.eecs.harvard.edu/~mdw/papers/codeblue-techrept05.pdf>.

[7]  	Vahed, Anwar.  “3-Lead Wireless ECG.”  15 November 2005.  29 September 2010. <http://www.open-ecg-project.org/tiki-download_file.php?fileId=5>.

[8]  	Vishay Lead (Pb)-Free Thick Film, Rectangular Chip Resistor Datasheet, at: www.vishay.com

[9]   Vishay Surface Mount Multilayer Ceramic Chip Capacitor for Commodity Applications   Datasheet, at: www.vishay.com

[10] Cornell Dubilier Electronics Type FCP Surface Mount Film Capacitor Datasheet, at: www.cde.com

[11] Burr-Brown Products fromTexas Instruments Single Supply CMOS Operational Amplifiers Zero-Drift Series Datasheet, at www.ti.com

[12] Kemet Ceramic Chip Capacitor Datasheet, at: www.kemet.com

[13] Panasonic Stacked Metallized Plastic Film Chip Capacitor Datasheet, at www.panasonic.com

[14] Raju, Murugavel.  “Heart-Rate and EKG Monitor Using the MSP430FG439.” October 2005. 12 September 2010.  <http://focus.ti.com/lit/an/slaa280a/slaa280a.pdf>.